| Date: | 23/05/2026 | Source Document No. | ERF/26/022 |
| --- | --- | --- | --- |
| Non-Conformance Description/ Description of event: |  |  |  |
| On 22/05/2026 during lyophilization (Freeze Drying) step of batch no. RIG25014, the vacuum pressure at the start of the primary drying step 1 was observed to be outside of specified limits between 20:59:11 to 21:06:11 (08 min elapsed, represented by eight one-minute data points). The lowest recorded vacuum pressure was at 192.40 µbar, against the acceptance range of 650 to 950 µbar which is critical process parameter (CPP). This observation was reported by Mr. Shrinath Yeotikar during routine batch monitoring. |  |  |  |
| Standard Procedures: |  |  |  |
| The vacuum pressure of lyophilizer (L-1901) shall be maintained within the range of 650 to 950 µbar, during the primary drying stage as it is critical process parameter (CPP) as per approved process control strategy report (GR/25/005 – R00) where vacuum pressure 1000 – 200 µbar and based on the ERF/25/146 – CAPA taken for change in pressure range i.e., as per below table. Step; Set Point (µ Bar); Proposed Controlling Band; Proposed Operating Range. Primary Drying – Step 1; 800; ±150; 650 – 950. Primary Drying – Step 2; 800; ±150; 650 – 950. Primary Drying – Step 3; 600; ±120; 480 – 720. Primary Drying – Step 4; 500; ±120; 380 – 620. Secondary Drying Step 1; 250; 200 – 600; 200 – 600. Secondary Drying Step 2; 250; 200 – 600; - – 600. |
| Immediate action taken (if any): |  |  |  |
| The observation was communicated to engineering team, Engineering team checked for the auto valves (YV-3106, YV-3107, YV-3108) to identify any malfunctioning, however no malfunction of valves observed. Further the manual vacuum control needle valve was adjusted to increase the pressure upon this action, pressure increased and maintained within the range, resulting in stabilization of the vacuum pressure and restoration of the system parameters within the acceptance range of 650 to 950 µbar which is critical process parameter (CPP) acceptable operating limits. |  |  |  |
| Investigation Team: |  |  |  |
| Name | Department (Role/Responsibility) |  |  |
| Sachin Kumbhar | Engineering -Team Lead |  |  |
| Gaurav Shelar | Engineering -Team member |  |  |
| Harshdeep Thakkar | Production – Team Member |  |  |
| Akash Kengar | Quality Assurance – Team Member |  |  |
| Investigation Tool Assigned: Documents review and Why-Why Analysis. |  |  |  |
| Chronology of the event: |  |  |  |
| Activity / Step; Observation / Details. Lyophilization cycle initiation; Lyophilization cycle for batch no. RIG25014 was initiated as per defined process parameters.. Shelf pre-cooling; Shelf pre-cooling activity was performed and shelf temperature of -5°C was achieved as per recipe requirements.. Freezing Control; Freezing stage was executed at -45°C with programmed ramp duration of 120 minutes and soak duration of 60 minutes.. Condenser pre-cooling; Condenser pre-cooling activity was performed and condenser temperature of -50°C was achieved with soak duration of 5 minutes.. Condenser pre-evacuation; Condenser pre-evacuation activity was performed with drying start pressure of 800.0 µbar. Alarm settings were maintained at P1: 900.0 µbar and P2: 950.0 µbar with P2 duration of 300 seconds as per recipe.. Primary Drying Step 1 initiation; Primary Drying Step 1 was initiated with set temperature of -45°C, ramp duration of 2 minutes, soak duration of 60 minutes, and chamber pressure of 800 µbar. Defined control band for pressure was 650 to 950 µbar.. Observation during Primary Drying Step 1; During routine monitoring between 20:59:11 to 21:06:11, chamber vacuum pressure was observed outside the specified control band and below the lower acceptance limit for approximately 08 minutes.. Lowest vacuum pressure observed; The lowest recorded chamber vacuum pressure observed during the event was 192.40 µbar against the specified control band of 650 to 950 µbar.. Reporting of observation; The observation was identified and reported by Mr. Srinath Yeotikar during routine batch monitoring activity.. Escalation of event; The observation was immediately communicated to the Quality team and Head of Production for further assessment and investigation.. |
| Investigation details: |  |  |  |
| - On 22/05/2026, during the execution of the primary drying (Step-1) in Lyophilizer (L-1901) for batch RIG25014, the vacuum pressure was recorded at 192.40 µbar, for 08 min, which is below the defined acceptance range of 650–950 µbar as per the BMR. - The engineering personnel, Mr. Gaurav Shelar acknowledged the observation at the same time. - Mr. Gaurav verified the nitrogen pressure regulator and found it to be satisfactory and to have approx. 1 bar, which is adequate for vacuum regulation. - The Gaurav Shelar inspected the automatic valves YV-3106, YV-3107, and YV-3108 to verify whether any malfunction could have restricted the nitrogen supply path and contributed to the sudden drop in vacuum. The inspection confirmed that all valves were functioning as intended, and no abnormalities or malfunctions were observed. - Thereafter, Mr. Gaurav Shelar verified the vacuum control needle valve. The needle valve was adjusted to increase the nitrogen supply for breaking the vacuum and restoring the chamber vacuum within the specified range (650–950 µbar). Due to the controlled and gradual adjustment of the nitrogen supply vacuum to stabilize within the specified operating range. - Refer attachment-2 Lyophilizer vacuum trend report. Figure-01 Lyophilizer Vacuum Trend (Batch no.: RIG25014) Process Understanding and Technical Evaluation - A detailed review of the vacuum-regulation mechanism of Lyophilizer L-1901 was conducted. - The investigation established that the vacuum-regulation system relies on manual adjustment of the vacuum-control needle valve and does not utilize an automated closed-loop feedback system for fine vacuum-pressure control. Consequently, chamber pressure control remains dependent upon manual intervention during process initiation and phase transitions. - The manual needle valve forms part of the nitrogen inlet and vacuum-control arrangement and provides controlled nitrogen bleed into the chamber to regulate vacuum pressure. - The investigation further confirmed that chamber pressure is not controlled by a single operating parameter. Rather, it is determined by the dynamic balance between: - Nitrogen inflow through needle valve ↔ Vapor generation/sublimation load ↔ Vacuum removal capacity - During primary drying, ice sublimates continuously from the product and generates water vapor inside the chamber. The sublimation rate is influenced by several variables including: - Product temperature - Shelf temperature - Residual ice content - Stage of lyophilization - Process transitions - Therefore, the vapor load entering the chamber is continuously changing throughout the freeze-drying cycle. - Simultaneously, nitrogen supply pressure is not an absolute constant variable. Even when the regulator setting remains unchanged, the actual pressure available across the needle valve may vary within the utility and control system. Since the needle valve functions as a throttling device, the amount of nitrogen entering the chamber depends on both valve opening and pressure differential across the valve. - As a result, the same physical needle-valve position may generate different nitrogen-flow rates under different process conditions. Process Condition; Sublimation/Vapor Load; Vacuum Tendency; Required Nitrogen Bleed. Higher sublimation rate; Higher vapor load; Chamber pressure tends to increase; Different valve setting required. Lower sublimation rate; Lower vapor load; Chamber pressure tends to decrease; Different valve setting required. Phase transition; Rapidly changing vapor load; Dynamic vacuum changes; Fine adjustment required. - Based on the above evaluation, it was established that a single fixed needle-valve position cannot maintain identical chamber pressure throughout the entire lyophilization cycle. - This is particularly relevant during process transitions and initiation of Primary Drying Step-1 because chamber vapor load changes rapidly while the process seeks a new equilibrium state. Detailed Investigation has been carried to determine root cause of excursion of vacuum out of specific range (650–950 µbar) - Post-batch completion. A detailed check was conducted by Mr. Gaurav on 24/05/2026 to identify contributing factors. - The nitrogen inlet line is equipped with a pressure regulator, solenoid valve (YV3011), and needle valve to control and stabilize vacuum pressure. - Nitrogen pressure regulator malfunction: Nitrogen supply pressure was observed at approximately 1 bar; this pressure was within the normal operating range of the nitrogen pressure regulator and was sufficient to provide the required nitrogen bleed for chamber vacuum regulation. Furthermore, after adjustment of the needle valve, the chamber vacuum stabilized and was maintained within the specified range of 650–950 µbar without any further intervention. This demonstrates that the available nitrogen supply pressure was adequate and did not contribute to the observed vacuum excursion. Hence, the cause of Nitrogen pressure regulator malfunction was ruled out. - Needle valve mechanical malfunction (sticking/jamming): Physical inspection and operational verification of needle valve performed. Needle valve movement was smooth without any sticking, blockage, wear, or damage. Hence, the cause of Needle valve mechanical malfunction was ruled out. - Needle valve inadvertent movement due to machine vibration: Visual inspection of valve position and machine operation performed, No evidence of valve displacement, loosening, or movement due to equipment vibration was observed. Valve remained in the adjusted position after stabilization. Hence, the cause of Needle valve inadvertent movement due to machine vibration was ruled out. - Vacuum pump malfunction or reduced pumping efficiency: Vacuum trend and equipment performance reviewed. The vacuum system achieved and maintained the required operating range after valve adjustment, indicating satisfactory pump performance. Hence, the cause of Vacuum pump malfunction ruled out. Refer attachment No.1 Lyophilizer Report - Chamber leakage or gasket failure: Vacuum stabilized and remained within specification after adjustment. No indication of continuous pressure rise or inability to maintain vacuum was observed. Hence, the cause of Chamber leakage or gasket failure ruled out. Refer attachment No.1 Lyophilizer Report - Vacuum sensor/transmitter malfunction: Instrument performance and pressure trend reviewed, Vacuum readings changed proportionally with valve adjustment and reflected actual process conditions also calibration status of vacuum transmitter has been verified and found within the acceptance criteria . No abnormal sensor behavior was observed. Hence, the cause of Vacuum sensor/transmitter malfunction is ruled out. Refer attachment No.2 Calibration Certificates of Vacuum Transmitter (PE-0251). - Manual adjustment sensitivity of needle valve: Event timeline and vacuum trend reviewed, Needle valve was observed in a partially open position. Due to the extremely sensitive nature of manual needle valve adjustment and the requirement for fine positioning, a stabilization period was required before achieving the target vacuum range. During this period, vacuum temporarily decreased to 192.40 µbar. Hence this Contributing factor identified. - Delay in achieving equilibrium after manual intervention: Process trend reviewed before and after adjustment. Following adjustment, vacuum gradually stabilized within the specified range and remained stable thereafter. This indicates a transient stabilization delay rather than equipment failure. Hence Contributing factor identified. Refer attachment No.03 Batch Trend Report of RIG 25014 - Operator-dependent manual control limitation: Vacuum regulation is performed manually through the needle valve without automatic feedback control. Minor changes in valve position significantly affect chamber pressure, making the process operator dependent. - The probable root cause of the event was the operator-dependent manual adjustment of the vacuum control needle valve. Also, the needle valve setting is dependent on the sublimation rate and the nitrogen supply pressure. Since both parameters vary during the lyophilization process, a fixed needle valve position cannot be established. - Due to the extremely sensitive nature of the needle valve and the time required for vacuum stabilization following fine manual adjustment, the chamber vacuum temporarily decreased below the specified range before equilibrium was established. No equipment or system malfunction was identified. - Also post completion of Batch No. RIG25014, vacuum trend was reviewed for same batch after primary drying stage, and no vacuum abnormalities were observed after primary drying till secondary drying or completion of batch as shown in fig.02 Fig. 02- Vacuum Trend- Batch No RIG25014 BMR Review A review of BMR-100-RIGC3-01 was performed to verify whether adequate procedural controls were available for vacuum regulation and needle-valve positioning during lyophilization. Review of Section 7.16.31 confirmed that specific instructions are provided for needle-valve positioning, wherein Production personnel are required to inform Engineering personnel, and Engineering personnel are required to verify and acknowledge the needle-valve position prior to execution of the relevant process step. The review confirmed that needle-valve positioning is an established and documented process control within the approved Batch Manufacturing Record and that the instruction was followed during execution of the batch. However, the existing instruction is limited to verification and acknowledgement of the needle-valve position and does not specifically address verification of vacuum-pressure stabilization during the transition into Primary Drying Step 1. Therefore, although the applicable BMR instruction was available and followed, a procedural gap was identified in that the current instruction does not adequately define the requirement to confirm vacuum-pressure stabilization following transition into Primary Drying Step 1. This gap may limit the effectiveness of the BMR in ensuring that the vacuum pressure is adequately stabilized before proceeding with the process step. Conclusion: The existing BMR instruction for needle-valve positioning was available and followed; however, a procedural gap was identified, as the current instruction does not adequately address confirmation of vacuum-pressure stabilization during the transition into Primary Drying Step 1. Accordingly, the BMR should be revised to incorporate appropriate instructions for verification and confirmation of vacuum-pressure stabilization during this critical process transition. SOP Review A review of applicable procedures was performed to determine whether adequate operational and maintenance controls were established for Lyophilizer L-1901 and its vacuum-control system. The following documents were reviewed: - SOP/P1/189 - SOP/P1/180 - SOP/P1/161 - SOP/EU/063 The review confirmed that approved procedures are available for operation, monitoring, maintenance and preventive maintenance of the lyophilizer and associated vacuum-control system. Records verified that preventive maintenance activities had been completed according to the approved maintenance planner and no overdue maintenance activity was identified that could have contributed to the observed event. The investigation further confirmed that operation of the equipment was performed in accordance with approved procedural requirements. Conclusion: Appropriate procedures were available and being followed. No procedural deficiency or maintenance-related issue was identified as a contributing factor. Training Review Training records of the Production and Engineering personnel concerned were reviewed to determine whether any training deficiency could have contributed to the observed vacuum excursion. The review confirmed that the personnel involved in operation, monitoring and verification of needle-valve positioning had completed the required training on applicable procedures and work instructions associated with lyophilizer operation and vacuum control. Training records were found complete, current and appropriately documented. Personnel acknowledgements were available and satisfactory. No evidence was identified indicating lack of knowledge, inadequate training, or failure to understand the applicable procedure. Conclusion: Training records were satisfactory, and no training-related deficiency was identified. Historical Review of Events: Sr. No; Event Description; Root Cause; Corrective Action; Preventive Action. 1; (ERF/25/135) On 24/10/2025 during lyophilization primary drying (step-1) for batch no. RIG25008 the vacuum pressure was observed to be outside of specified limits between 07:38 to 07:53 (15 min.). The recorded vacuum pressure was in 300 µbar, against the BMR acceptance range of 500 to 1000 µbar.; On 24/10/2025 during lyophilization primary drying (step-1) for batch no. RIG25008 the vacuum pressure was observed to be outside of specified limits between 07:38 to 07:53 (15 min.). The recorded vacuum pressure was in 300 µbar, against the BMR acceptance range of 500 to 1000 µbar.; The needle valve was manually adjusted based on the work order during the batch process to restore and maintain the vacuum pressure at 800 µbar, ensuring the pressure remained within the specified range as required. As predictive maintenance, needle valve replaced with a new one.; A revised BMR step will be incorporating requiring: 1.After adjustment of needle valve by engineering personnel, mandatory verification and documentation of needle valve shall be performed prior to starting primary drying supported by operator and supervisor dual sign off. 2; (ERF/25/146) On 05/11/2025, during the lyophilization primary drying stage (Step-4) of Recombinant Insulin Glargine Batch No. RIG25009, the vacuum pressure was observed to be outside the specified momentary fluctuation limits between 09:38 and 15:07 hours (approximately 5 hours and 31 minutes). The recorded vacuum pressure ranged from a minimum of 447 µbar to a maximum of 545 µbar, against the BMR-RIGC3-00 acceptance criterion of 500–1000 µbar (CPP) during the primary drying stage (Step-4).; The observed vacuum pressure excursion during the primary drying stage (Step-4) of Batch No. RIG25009 attributed to the inherent control system bandwidth characteristics. The vacuum control system designed to maintain pressure around the set point; however, due to its control bandwidth and response characteristics, it permits temporary deviations below the set point. In this instance, although the set point was 500 µbar, the system allowed fluctuations down to 447 µbar. Since the defined acceptance range for primary drying is 500–1000 µbar, these inherent system fluctuations resulted in a momentary deviation outside the specified limits. The root cause can be categorized “Machine / Equipment.”; Based on the performance across executed batches, an allowance range for vacuum pressure will be formally defined for each step of the freeze-drying process. The revised controlling bands and operating ranges will be implemented in the process control documentation and used for monitoring and evaluation during ongoing and future cycles. Step; Set Point (µ Bar); Proposed Controlling Band; Proposed Operating Range. Primary Drying – Step 1; 800; ±150; 650 – 950. Primary Drying – Step 2/3; 600; ±120; 480 – 720. Primary Drying – Step 4; 500; ±120; 380 – 620. Secondary Drying Step 1 /2; 250; 200-600; 200-600. These changes shall be incorporated into BMR-100-RIGC3. The data suggests that there are momentary excursions reported during cycle are inherent to the system control. Hence this note shall be included in BMR-100-RIGC3. Additionally, it is proposed to revise the vacuum pressure range currently specified in the BMR from a single value of 500–1000 µbar to stepwise set points for each stage of primary and secondary drying, ensuring alignment with the actual recipe of the Lyophilization cycle. This shall be also incorporated in BMR-100-RIGC3 and Process Control Strategy document.; NA. 3; (ERF/25/186) On 07/02/2026, during the batch processing of Recombinant Human Insulin Batch No. RHI25063, it was observed that during batch processing in the Lyophilizer, the vacuum pressure exceeded the defined upper limit during the primary drying stage (step 04). At approximately 19:45 Hrs., the vacuum pressure was showed 474.80 micro bar against a set point of 500 micro bar. However, at approximately 20:02 Hrs., the vacuum pressure increased to 874.40 micro bar, which is beyond the defined limit.; Based on the detailed investigation and the outcome, Intermittent malfunction of refrigeration expansion valve causing restricted refrigerant flow, resulting in reduced condenser cooling capacity, elevated condenser temperature, and subsequent vacuum pressure excursion during primary drying. Also It was revealed that the Batch Manufacturing Record (BMR) did not include a procedure to verify.; - The needle valve was manually adjusted based on the work order during the batch process to restore and maintain the vacuum pressure at 800 µbar, ensuring the pressure remained within the specified range as required. - As predictive maintenance, expansion valve replaced with a new one.; Historical review of vacuum pressure performance of previously executed batches will be conducted and based on this evaluation, appropriate operating ranges and control bands for each freeze-drying step will be defined and justified. The finalized ranges will be incorporated in the process BMR and used for monitoring & assessment.. Review of Historical Events and CAPA Effectiveness Assessment - A review of previous vacuum pressure-related deviations associated with Lyophilizer L-1901 was performed. - ERF/25/135 was attributed to improper needle valve positioning, for which corrective actions included manual adjustment of the needle valve and replacement of the needle valve as a predictive maintenance activity. Preventive actions included implementation of mandatory verification and documentation of needle valve position prior to initiation of primary drying, supported by operator and supervisor dual sign-off. - ERF/25/146 was attributed to inherent vacuum control system fluctuations. As a result, stepwise vacuum operating ranges and controlling bands were established based on historical process performance and incorporated into the Batch Manufacturing Record (BMR) . - ERF/25/186 was attributed primarily to intermittent malfunction of the refrigeration expansion valve resulting in condenser performance deterioration and subsequent vacuum pressure excursion. In addition, preventive actions included mandatory verification and documentation of needle valve position before initiation of primary drying, operator-supervisor dual sign-off, retraining of operating personnel on lyophilizer vacuum control parameters. - Verification performed during the current investigation confirmed that all CAPA actions arising from the above historical events had been implemented and were being followed and personnel had been trained as required. Furthermore, no evidence of refrigeration system malfunction, expansion valve failure, vacuum pump deficiency, nitrogen supply issue, instrumentation failure, chamber leakage, or needle valve mechanical defect was identified during the current event. Inference: A review of historical vacuum-pressure excursions associated with Lyophilizer L-1901 was performed to evaluate recurrence trends and effectiveness of previously implemented CAPAs. - The review confirmed that the root causes identified in historical events ERF/25/135, ERF/25/146 and ERF/25/186 were different in nature and were addressed through specific corrective and preventive actions. These actions included needle-valve replacement, implementation of procedural controls for needle-valve verification, dual sign-off requirements, personnel retraining, establishment of operating ranges and controlling bands, refrigeration-system repairs, and preventive maintenance activities. - Verification conducted during the current investigation confirmed that all previously approved CAPAs were implemented and sustained. Furthermore, the present investigation found no evidence of recurrence of the specific root causes identified in the historical events. The needle valve was functioning properly, refrigeration performance was satisfactory, nitrogen supply was adequate, instrumentation was functioning normally, and established procedural controls were being followed. - Although the previously implemented CAPAs were effective in addressing their respective root causes and reducing the likelihood of recurrence, they could not completely eliminate the residual risk associated with the existing vacuum-control configuration of Lyophilizer L-1901. The system continues to rely on manual needle-valve adjustment for fine vacuum regulation because the installed vacuum-control arrangement does not provide automated closed-loop modulation capable of maintaining the chamber pressure at the desired setpoint under all process conditions. - Historical batch-trend review demonstrated that transient vacuum excursions are observed predominantly during phase transitions and vacuum-stabilization periods. The occurrence of these events is influenced by process dynamics, including variations in sublimation rate, vapor load, chamber-pressure response and manual needle-valve adjustment requirements. Consequently, the required needle-valve position may vary throughout the lyophilization cycle and a single fixed valve position cannot ensure identical chamber-pressure control under all operating conditions. - The current investigation established that the observed vacuum excursion was associated with the dynamic response of the vacuum-control system during manual needle-valve adjustment and subsequent stabilization at the initiation of Primary Drying Step-1. Therefore, the event is considered related to the inherent operational limitation of the existing vacuum-control strategy rather than failure of equipment, utilities, instrumentation, procedures, training or previously implemented CAPAs. - Accordingly, the CAPAs from the historical events are considered effective for controlling the specific causes for which they were implemented; however, they did not fully eliminate the residual process risk associated with manual vacuum regulation and stabilization under varying lyophilization conditions. Historical Review of Batches : Batch; Step; Setpoint; Specified Range; Observed Range; Excursion Detail; Start and end Date/Time; Total duration of vacuum excursion (min); Direction; Remark. RIG23001; Primary Drying – Step 1; 800; 500–1000 µbar; 1000 Max; 13 readings at Max.1000 to 934 µbar; 22/03/2024 04:06:01 to 04:19:01; 13 Minutes; High; At transition into Step 1 -800 µbar setpoint. RIG23008; Primary Drying – Step 1; 800; 500–1000 µbar; 201.03 Min; 16 readings at 201.03 Min; 20/04/2024 18:55:02 to 19:11:02; 16 Minutes; Low; At transition into Step 1000 -800µ bar setpoint. RIG24003; Primary Drying – Step 1; 800; 500–1000 µbar; 502.50 Min; 10 reading at Min 502.50; 11/05/2024 04:39:44 to 04:49:44; 10 Minutes; Low; At transition into Step 1000-800µ bar setpoint. RIG25010; Primary Drying – Step 4; 500; 380–620; 325.9 min; 1 reading at 325.9 µbar; 18/11/2025 05:37:32; 1 Minute; Low; Transient low reading; subsequent values within range.. RIG25014; Primary Drying – Step 1; 800; 650–950; 192.4 min; 8 consecutive one-minute readings outside range; 22/05/2026 20:59:11–21:06:11; 8Minutes; Low; Corresponds to ERF/26/022; investigation attributes events to sensitive manual needle-valve control and stabilization.. C072630015; Primary Drying – Step 3; 600; 480–720; 755.3 max; 24 one-minute readings above 720; 29/05/2026 22:27:27–23:38:27; 1 minutes readings at different time; High; Intermittent high-side excursions.. C072630017; Primary Drying – Step 4; 500; 380–620; 631.3 max; 1 reading at 631.3; 08/06/2026 11:27:08; 1 Minute; High; At transition into Step 4.. Inference: The batch comparison demonstrates that transient vacuum excursions are not uniformly observed in every batch or every phase. Several excursions occur at or immediately after phase transitions (Freezing step to Primary drying Step 1) while some batches remain fully within the proposed operating ranges. These supports treating manual needle-valve adjustment and subsequent vacuum equilibration as a controlled process activity. The ERF/26/022 investigation specifically concludes that the Lyophilizer vacuum regulation system relies on manual needle-valve adjustment, that the valve is sensitive, and that stabilization time is required following fine manual adjustment. The investigation proposes documenting a 20-minute period for needle-valve adjustment/stabilization. The lyophilizer vacuum control system consists of an EV valve, which operates only in ON/OFF mode and therefore does not provide precise modulation required to maintain micro-level vacuum at the defined setpoint. A manually adjustable needle valve is provided for finer vacuum regulation and to minimize fluctuations. However, due to the inherent limitations of the ON/OFF EV valve and the dependence on manual needle valve adjustment, transient vacuum fluctuations may occur. In addition, the sublimation rate varies across different drying steps, resulting in changes in vapor load and corresponding vacuum fluctuations. Considering this equipment and process-related operational limitations, a 20-minute duration is considered an appropriate and practical allowance for stabilization of vacuum following a transient excursion. This duration provides sufficient time for the vacuum to recover within the defined operating range without requiring unnecessary operator intervention, while remaining consistent with the historical operating behavior of the lyophilizer. The historical trend review, together with the current investigation, indicates that the observed vacuum excursions are associated with the dynamic response of the vacuum-control system following manual needle-valve intervention during process transition in Primary Drying Step 1, rather than recurrence due to failure of previously implemented CAPAs. The previously implemented CAPAs were reviewed and confirmed to have been implemented as committed and were appropriate for addressing the specific procedural and preventive control gaps identified at the time of the respective events. However, the recurrence of a pressure excursion during the transition into Primary Drying Step 1 demonstrates that these actions have not fully controlled the underlying vacuum-control risk associated with manual needle-valve operation and vacuum stabilization. In particular, the existing control of verifying the needle-valve position provides assurance regarding the valve setting but does not ensure stable and controlled chamber pressure during the transition into Primary Drying Step 1. The current event therefore indicates that the residual risk associated with manual vacuum adjustment remains, particularly due to the sensitivity of the needle valve and the requirement for fine manual adjustment during pressure stabilization. Accordingly, the previously implemented CAPAs are considered implemented and effective in addressing the specific deficiencies for which they were established; however, they are not considered fully effective in eliminating or adequately mitigating the risk associated with manual vacuum control during the transition into Primary Drying Step 1. - Photograph of Precise Needle Valve and Proposed Control: Proposed Controls: Short term Actions: 1. Preparation and execution of Lyophilizer Vacuum Control Study Protocol A formal Lyophilizer Vacuum Control and Needle-Valve Position Study Protocol shall be prepared and executed to establish the optimum needle-valve position and vacuum-control response for each applicable lyophilization step. The study shall evaluate the vacuum response during transition and stabilization for the following approved operating conditions: Lyophilization Step; Set Point (µbar); Controlling Band (µbar); Operating Range (µbar). Primary Drying – Step 1; 800; ±150; 650–950. Primary Drying – Step 2; 800; ±150; 650–950. Primary Drying – Step 3; 600; ±120; 480–720. Primary Drying – Step 4; 500; ±120; 380–620. Secondary Drying – Step 1; 250; 200–600; 200–600. Secondary Drying – Step 2; 250; 200–600; 200–600. 2. Determination of optimum initial needle-valve position During the study, the initial needle-valve position for each lyophilization step shall be established and documented. The study shall assess, as applicable: - Initial needle-valve position before initiation of each step. - Needle-valve position corresponding to the respective vacuum set point. - Vacuum response following set-point transition. - Direction and magnitude of needle-valve adjustment required. - Repeatability of the identified valve position over multiple cycles Same shall be established and incorporated into the applicable One Page Lesson The objective shall be to establish a defined and reproducible starting position for the needle valve for each recipe step, thereby minimizing operator-dependent adjustment. 3. Vacuum alarm challenge test A Vacuum Alarm Challenge Test shall be included in the study protocol to verify the functionality of the lyophilizer's visual and audible vacuum alarms against the step-specific operating ranges. The challenge shall verify, for each applicable lyophilization step: - Alarm indication when chamber vacuum exceeds the defined operating range. - Visual alarm indication. - Audible alarm indication. 4. Training Module A dedicated Lyophilizer Vacuum Control and Needle-Valve Operation Training Module shall be prepared based on the approved study protocol and study outcome. The training shall cover: - Lyophilizer recipe and step sequence. - Step-specific vacuum set points and operating ranges. - Initial needle-valve positions. - Relationship between needle-valve movement and chamber vacuum. - Correct incremental adjustment technique. - Vacuum trend interpretation. - Alarm indication Training shall be imparted to all concerned personnel involved in lyophilizer operation, monitoring, supervision and troubleshooting, and training effectiveness shall be documented as per the applicable training procedure. 5. One Page Lesson A One Page Lesson shall be prepared and displayed/available at the applicable operator workstation. Investigation done with Why-Why analysis: Why-Why Analysis Q1. Why did the chamber vacuum decrease to 192.40 µbar during Primary Drying Step-1? A1. Because the nitrogen bleed into the chamber was insufficient to counterbalance the vacuum being generated, resulting in the chamber pressure decreasing below the approved operating range of 650-950 µbar. Q2. Why was the nitrogen bleed into the chamber insufficient? A2. The nitrogen bleed flow to the chamber is not directly measured by a flow transmitter or flowmeter; therefore, the actual flow during the event cannot be quantitatively confirmed. The effectiveness of the nitrogen bleed was assessed based on the chamber pressure response following needle-valve adjustment, along with the available needle-valve position and nitrogen supply-pressure information. The manually operated needle valve controls the nitrogen bleed into the chamber. Its opening directly affects the nitrogen flow, with smaller openings providing lower flow and larger openings providing higher flow, depending on the nitrogen supply pressure and process conditions. During the transition into Primary Drying Step 1, vacuum demand may vary with process conditions. Therefore, the same needle-valve position may not always maintain the required chamber pressure, resulting in temporary pressure fluctuations until the valve is adjusted and pressure is stabilized . Q3. Why did the needle-valve opening not provide the required nitrogen flow under the prevailing process conditions? A3. Because during initiation of Primary Drying Step-1, the process conditions were changing and the chamber vacuum required further fine adjustment of the needle valve before a new pressure equilibrium could be established. Q4. Why was further fine adjustment required after the needle-valve position had already been verified? A4. Because the required needle-valve position is dependent upon chamber-process conditions, including sublimation rate, vapor load and nitrogen-flow demand. These parameters change throughout the lyophilization cycle, particularly during phase transitions. As a result, one fixed needle-valve position cannot consistently maintain the target chamber pressure under all operating conditions. Q5. Why does the process depend on further manual adjustment when process conditions change? A5. Because the Lyophilizer L-1901 vacuum-control system relies on manual adjustment of the vacuum-control needle valve for fine vacuum regulation. The installed EV valve functions primarily in an ON/OFF mode and does not provide continuous closed-loop modulation capable of automatically maintaining chamber pressure at the desired setpoint under varying process conditions. Q6. Why did the event occur despite implementation of previous CAPAs and compliance with approved procedures? A6. Because the CAPAs implemented from previous events effectively addressed the specific causes identified at that time, including needle-valve replacement, procedural verification, dual sign-off, operating-range definition, preventive maintenance and personnel training. However, these CAPAs did not eliminate the inherent dependence of the vacuum-control system on manual needle-valve adjustment and the associated stabilization period required during changing process conditions. Therefore, a residual possibility of temporary vacuum excursion remains despite compliance with established controls. Q7. Why could the residual risk not be completely eliminated through the existing control strategy? A7. Because the current vacuum-control configuration inherently depends on manual regulation of the needle valve in response to changing chamber conditions. The required valve position varies with sublimation rate, vapor generation, nitrogen-flow requirements and vacuum response during the lyophilization cycle. Consequently, the implemented controls reduce the likelihood and impact of recurrence but cannot completely eliminate the process variability associated with manual vacuum regulation and stabilization. |
| Root Cause/ Probable Cause: |  |  |  |
| The root cause of the event was the vacuum-control system limitation of Lyophilizer L-1901, which relies on a manually operated needle valve for fine chamber-pressure adjustment during the transition into Primary Drying Step 1. The installed system does not provide continuous closed-loop fine pressure modulation to automatically compensate for changes in vacuum demand and process conditions. Consequently, manual needle-valve adjustment is required to establish and stabilize the chamber pressure within the specified operating range. The investigation confirmed that the operators were trained and followed the approved procedure, and no mechanical failure of the needle valve was identified. Therefore, the event is not attributed to operator error or performance, but to the inherent limitation of the existing vacuum-control system and its dependence on manual intervention for pressure stabilization. Due to changes in sublimation rate, vapor load and chamber-pressure response during process transition, the required needle-valve setting could not be represented by a single fixed position. Consequently, additional fine adjustment and vacuum-stabilization time were required to achieve equilibrium. Contributing Factors - High sensitivity of the vacuum-control needle valve. - Absence of automated closed-loop chamber-pressure control. - Dynamic changes in sublimation rate and vapor load during process transitions. - Time required for chamber-vacuum stabilization following manual adjustment. - Residual process variability associated with manual vacuum regulation. |  |  |  |
| Impact Assessment: |  |  |  |
| The observed excursion of the critical process parameter beyond the control band occurred only during the initiation of the primary drying stage and was transient in nature. The deviation was attributed to the sensitivity of the manually operated needle valve during throttling adjustments and not to any equipment malfunction or process failure. The critical process parameter (CPP) was subsequently stabilized within the prescribed operating range through continued monitoring and adjustment. The additional contributing factors, namely the requirement for fine-tuning due to valve sensitivity, increased the likelihood of operator-dependent variation during vacuum stabilization. The shelf temperature and product temperature trends during the pressure excursion and subsequent recovery were reviewed and found to be within the established process operating ranges, with no abnormal temperature trend observed. The reviewed trends are also attached in the BMR and Attachment No.03 for Investigation Report The impact of the transient lower chamber pressure on sublimation and drying was also assessed. No abnormal change in product temperature or process behavior indicative of an adverse effect on sublimation was observed during the excursion or recovery. The primary drying process continued as expected, with no evidence of affected or uncontrolled sublimation or abnormal drying behavior. The HMWP result (0.21%), LOD result (4.09%), and Assay result (101.3%) were within the specified acceptance criteria. In particular, the LOD result supports adequate drying, while the HMWP and Assay results provide additional assurance that no adverse impact on product quality was identified. Based on the combined assessment of chamber pressure behavior, shelf and product temperature trends, process observations, and available analytical results, no adverse impact on the sublimation/drying process or product quality was identified due to the transient pressure excursion. ., the table below is the final Certificate of Analysis (COA) Test; Specification Limits; Result. Appearance; White or almost white powder.; White powder. Identification. Identification A: Assay; The retention time of the major peak of the Sample solution corresponds to that of the Standard solution, as obtained in the Assay.; Complies. Identification B: Peptide Mapping; The chromatographic profile of the Sample solution corresponds to that of the Standard solution. All four fragments, I, II, III and IV must be present.; Complies. Limit of High Molecular Weight Proteins; NMT 0.3%.; 0.21%. Product Related Substances and Impurities. Any Individual Insulin Glargine related substance; NMT 0.5%; 0.08. Total Insulin Glargine related substances; NMT 1.5%; 0.31. Zinc Determination; NMT 0.80%; 0.3. Loss on Drying; NMT 10.0%; 4.09. Bacterial Endotoxins test; NMT 10 EU/mg; <10 EU/mg. Assay; 94.0% - 105.0% on dried basis.; 101.3. Host Cell Proteins; NMT 10 ppm (ng/mg).; 2. Microbial Limit Test. Microbial Enumeration tests; NMT 300 cfu /g; 25 cfu /g. Test for specific Microorganisms: E. coli, Salmonella, P. aeruginosa, S. aureus, Clostridium sporogenes, Bile tolerant gram-negative bacteria, Candida albicans.; Absent/g; Absent. *Host Cell DNA; NMT 1.3 ng/mg (10 ng/dose); <1.3 ng/mg (10 ng/dose). *Residual Solvents. Acetonitrile; NMT 410 ppm.; 9 ppm. Isopropyl Alcohol; NMT 5000 ppm.; 24 ppm. Note: *Additional tests. The analytical results for final drug substance batch RIG25014 were reviewed. Based on the completed analytical testing available to date, all reported results were within their specified acceptance limits, and no adverse impact on product quality has been identified. Microbiological and endotoxin results were also within specified limits, supporting product safety and purity. |
| Scope Assessment: |  |  |  |
| The evaluation confirmed that the specific deviation investigated was observed in batch RIG25014. Review of concurrent batch data confirmed that no comparable pressure excursion was observed during the Primary Drying Step 1 transition for batches C072630015, C072630016, and C072630017. However, the historical review identified similar pressure excursions during the transition into Primary Drying Step 1 in other batches. Based on the historical trend, the observed pressure excursion in RIG25014 is not considered an isolated batch-specific event. Rather, the recurrence of similar pressure excursions indicates that the event is associated with the existing vacuum-control system behavior during the transition into Primary Drying Step 1. Therefore, RIG25014 represents a specific occurrence of a recurring vacuum-control behavior, rather than an isolated deviation attributable solely to the batch |  |  |  |
| Corrective Action: |  |  |  |
| Restoration of Vacuum Pressure During Batch Processing During batch processing, the vacuum-control needle valve was manually adjusted by Engineering personnel based on process feedback and communication received from the Production team. The adjustment was performed to increase the nitrogen bleed into the chamber and restore the vacuum pressure to approximately 800 µbar, thereby ensuring that the chamber vacuum returned to and remained within the specified operating range as per process requirements. |  |  |  |
| Preventive Action: |  |  |  |
| Short Term Actions : 1. Preparation and execution of Lyophilizer Vacuum Control Study Protocol ,Vacuum alarm challenge test ,Training Module,One page Lesson as per reference in attachment No.4 . 2. Revision of BMR to Define Vacuum Stabilization Requirements BMR-100-RIGC3-01, Section 7.16.31, shall be revised to include a defined 20-minute vacuum stabilization period following manual needle-valve adjustment during Primary Drying Step-1. and monitoring of chamber-vacuum trends during the stabilization period to verify recovery of vacuum within the approved operating range. This revision is based on historical batch-trend review, equipment operating characteristics, and the findings of Investigation ERF/26/022, which demonstrated that vacuum stabilization following fine manual needle-valve adjustment may require a defined equilibration period before steady-state operating conditions are achieved. If vacuum pressure does not stabilize within the specified period, the applicable event escalation procedure shall be initiated. This action is intended to strengthen procedural controls and mitigate the residual risk associated with manual vacuum regulation and stabilization identified during the investigation. 3.Update of FMEA, Process Control Strategy (PCS) . As an outcome of Investigation ERF/26/022, the applicable FMEA and PCS shall be revised through the approved change-control/CAPA system. The FMEA shall be updated to include the revised preventive and detection controls, including reference to BMR-100-RIGC3-01, monitoring of chamber-vacuum trends, a defined 20-minute stabilization period following manual needle-valve adjustment during Primary Drying Step-1, and initiation of the applicable QMS event-escalation procedure if the vacuum does not stabilize within the approved operating range. The occurrence and detectability ratings shall be reassessed based on the implemented controls and CAPA effectiveness verification. Additionally, the existing generalized vacuum range of 1000–200 µbar in PCS Table No. 4 shall be replaced with the following step-specific set points, controlling bands, and operating ranges: Step; Set Point (µ Bar); Proposed Controlling Band; Proposed Operating Range. Primary Drying – Step 1; 800; ±150; 650 – 950. Primary Drying – Step 2; 800; ±150; 650 – 950. Primary Drying – Step 3; 600; ±120; 480 – 720. Primary Drying – Step 4; 500; ±120; 380 – 620. Secondary Drying Step 1; 250; 200 – 600; 200 – 600. Secondary Drying Step 2; 250; 200 – 600; - – 600. Note: Following manual adjustment of the vacuum-control needle valve during Primary Drying Step-1, the chamber vacuum shall be monitored for a stabilization period of 20 minutes. If the vacuum does not stabilize within the approved operating range during this period, the applicable QMS event-escalation procedure shall be initiated. Based on the investigation findings and historical batch-trend review, the observed vacuum variation did not adversely affect the evaluated product quality attributes or process performance. Therefore, the residual risk associated with the current manual needle-valve vacuum-control system is considered acceptable, subject to implementation and effectiveness verification of the revised controls. Long Term Action : 1. Evaluation and Implementation of Automated Vacuum Control System An engineering evaluation shall be performed to assess the feasibility of implementing a highly precise automated MKS PID control valve for regulating chamber vacuum during all stages of the lyophilization cycle and The automated vacuum control system considered a long term actionThe proposed automated control system will provide continuous and accurate vacuum regulation, thereby reducing dependence on manual needle-valve adjustment. Implementation of this control strategy is expected to minimize vacuum fluctuations associated with manual intervention, improve process stability, enhance compliance with approved operating limits, and improve overall process reliability and reproducibility. |
| List of Attachment: |  |  |  |
| - Batch Trend reports of RIG25009, RIG25010, RIG25013, RIG25014, C072630015, C072630015, C072630016, C072630017. - Calibration certificates of Vaccum Transmitter (PE-0251) - Batch trends of Shelf temperature and Product temperature of RIG 25014. - Justification Report for 20 minute vacuum stabilization/Needle valve adjustment and Proposed Controls. |  |  |  |

| Prepared By EU (Sign. & Date) | Reviewed By EU (Sign. &Date) | Reviewed By Production (Sign. & Date) | Reviewed By QA (Sign. & Date) | Approved By QA (Sign. &Date) |
| --- | --- | --- | --- | --- |