# 21 CFR Part 11 & ALCOA++ Design Specification Matrix

Purpose:

This matrix consolidates 21 CFR Part 11 electronic record/electronic signature requirements with ALCOA++ data integrity attributes into a single set of software design specifications. It is intended for use as a baseline requirement reference when evaluating or specifying GxP software from a vendor. Procedural/organisational-only clauses (e.g., personnel training, signature accountability policy) are excluded as they are not software design features.

| 21 CFR Part 11 Clause | ALCOA++ Attribute(s) Addressed | Required Software Feature / Design Specification |
| --- | --- | --- |
| §11.10(a) | Foundational – underpins all attributes | System shall be validated via documented IQ/OQ/PQ with risk-based test coverage across all GxP functions before production use. |
| §11.10(b) | Legible, Original, Available | System shall export the complete record, full audit trail, and associated metadata in both human-readable (PDF/print) and electronic (XML/CSV) form, with no loss of context versus the source record. |
| §11.10(c) | Enduring, Available | System shall provide a defined data retention and archival mechanism with periodic restore verification, and indexed retrieval of archived records within an agreed SLA. |
| §11.10(d) | Attributable | System shall enforce role-based access control with unique individual user accounts. Shared or generic logins shall not be permitted. Access revocation shall be supported. |
| §11.10(e) | Attributable, Contemporaneous, Accurate, Complete | System shall auto-generate a secure, non-editable audit trail for every create, modify, and delete action on GxP data, capturing old value, new value, user ID, system-generated timestamp, and reason for change. Original data shall never be overwritten. All data, including aborted or repeated events, shall be retained. |
| §11.10(f) | Consistent | System shall enforce permitted sequencing of process steps via workflow/state-machine logic, preventing out-of-sequence actions (e.g., approval prior to review). |
| §11.10(g) | Attributable | System shall perform authority checks tied to both role and task, and shall support segregation of duties (e.g., reviewer cannot also act as approver for the same record). |
| §11.10(h) | Consistent | Where data originates from interfaced instruments/devices, system shall validate the identity/source of the originating device prior to data acceptance. |
| §11.30 | Supports Original, Accurate (open systems) | For open-system architectures, system shall provide encryption of data in transit and at rest. |
| §11.50 | Attributable | Electronic signature manifestation shall display the signer's printed name, date/time of signing, and the meaning of the signature (e.g., review, approval, authorship). |
| §11.70 | Attributable, Original | Electronic signatures shall be logically/cryptographically bound to the specific record version signed. The system shall prevent copying or transferring a signature to falsify another record. |
| §11.100 | Attributable | System shall enforce a unique user ID per individual at account creation. Deactivated user IDs shall not be reused or reassigned. |
| §11.200 | Attributable | Electronic signature execution shall require two distinct components (ID code + password) or biometric verification, with re-authentication required at the moment of signing – not satisfied by an active login session alone. |
| §11.300 | Attributable | System shall enforce password complexity and expiry, account lockout after a defined number of failed attempts, and a deactivation workflow for lost or compromised credentials. |