| Date: | 15/04/2026 | Deviation No. | DEV/IT/25/002 |
| --- | --- | --- | --- |
| Investigation tool used: 6M 5 Why Brainstorming |  |  |  |
| Other Tools (If any): |  |  |  |
| Not Applicable |  |  |  |
| Define: |  |  |  |
| Following checks shall be considered while writing the “Define” section. - Clearly define what happens actually. - Explain what is different than expected. - Mention the location at where the deviation has occurred. - Date/time of deviation occurrence and date/time of detection. - Mention the name of Personnel who is involved in the deviation. - Mention initial scope of deviation (i.e., impacted product/Material/Equipment/System//Batches/etc.) On dated 17/11/2025 QA Personnel (Priya Patole-584 ) review the audit trail in QC Department, and it is observed that the audit trail data was not available of PH Conductivity meter (I/QC/10) for 14/10/2025 to 28/10/2025. As per the SOP/DP/QA/003, data shall be available throughout the period of document retention period and as per SOP/DP/IT/001, instrument procedure data shall be available in the system for Audit Trail. After this incident QA Personnel inform to IT and QC concerned Personnel. IT Personnel visited the system and data were not available in system. IT Personnel copied latest available backup from NAS Server and checked that from 14/10/2005 to 27/10/2025 data is retrieved, however, for 28/10/2025 data not available in system. vendor visited on the site 05/02/2026 to assess the issue and explore the possibility of data retrieval. But for 28/10/2025 data could not be restored from the available backup with IT due to system overwritten functionality. Vendor provided conclusion report on 21/03/2026 at 17:12. Therefore it is decided to take deviation for which data could not be retrieved. |  |  |  |
| Details Investigation: |  |  |  |
| Measure: Following checks shall be considered while writing the “Measure” section. - Does the summary provide relevant facts and data/information that was reviewed including: environment, process/product history, Personnel info (title and job title), controls/control limits, etc. - Is a summary of the analysis of the factors and data provided? - Is a conclusion statement of the analysis and review provided? - If there were Regulatory Notification, were details provided? - Is the report written in a logical flow and easily understood by the reader? On dated 17/11/2025 QA Personnel (Priya Patole-584 ) review the audit trail in QC Department, and it is observed that the audit trail data was not available of PH Conductivity meter (I/QC/010) for 14/10/2025 to 28/10/2025. As per the instrument procedure data shall be available in the system for Audit Trail. As per the QA SOP No. SOP/DP/QA/029/F02-R00 QA Personnel is execute the audit trail activity SOP/DP/QA/029/F02-R00(Audit Trail Review Checklist). After the audit trail review need to fill details in Audit Trail Review Sheet as per Format No. SOP/DP/QA/029/F01-R00. According to that QA Personnel visited the instrument I/QC/010 on dated 17/11/2025. QA Personnel is mentioned the list of equipment /instrument/system with audit trail for each department as per format-III (SOP/DP/QA/029/F03). QA is reviewing the data for two types of GxP System one is Direct impacted like filling machine, HPLC, Vessels, LBPC and second is Indirect impacted systems like water system, vitek, incubators. QA audit trail review frequency is as per requirement before batch or every month within 7 days for direct impacting and indirect impacting GxP Systems. QA is Verifying changes in parameters, set points, privilege matrix, policies, and configurations. Ensure audit trails are enabled and data stored at the correct path. QA is Checking for unapproved changes, invalid logins, unauthorized time/date edits, duplicate entries/data, and deletions or edits of records. Confirm that any duplicate entries are properly approved. Any Document discrepancies QA communicate to department heads. Ensure corrective actions are taken and recorded. Once all points verified by QA is ensuring that all points meet as per criteria if any discrepancy found QA must inform CFT and QA HOD for further procedure. As per the scheduled QA Personnel reviewed data for I/QC/010 Instrument and observed that the audit trail data has been missing from 14/10/2025 to 28/10/2025. After that QA Personnel informed concerned department Head and IT Personnel. IT Personnel visited the system and checked data availability on server for mentioned dates. Data backup was available with IT; IT Personnel copies that data backup and stored it in that system data backup path location. After the data backup was stored, QA Personnel checked that it has recovered from 14/10/2025 to 27/10/2025. Audit trail data was missing for date 28/10/2025. IT Personnel and QC Personnel contacted Service Engineer to visit the site for the missing data rectification. On dated 05/02/2025 Service Engineer visited on MJ Biopharm Site for data missing rectification. Service Engineer tried to all possibilities to data recover from available backup, but same issue persists. After several attempts it is observed that data is not capturing in data backup file for 28/10/2025 hence it has unable to restore as required. Vendor spoke with his atomization team, and at the end it concluded that system having overwritten functionality of backup creation file. If you are taken backup and after that file need to be renamed otherwise it is overwritten on that previous backup file. However, IT Personnel taken regular backup on dated 27/10/2025 and copy that file on server, then after system memory get full on dated 29/10/2025 hence IT Personnel taken Backup on 29/10/2025 before archival and QA Personnel archived data for free space in system on same day. Which backup taken on 29/10/2025 is having full backup till 28/10/2025 however when IT taken backup on 29/10/2025 is not renamed backup file, therefore monthly backup in next month taken that time previous file is overwritten by system and new backup file is generated. Hence 28/10/2025 data was not available in that backup file which led to missing data for mentioned data missing date. The batch data has been available, hence no impact is anticipated, however regulatory notification is not required. |  |  |  |
| Analyze: Various investigation tools used to analyze the root cause like 6M, 5 Why, Personnel interview, Brainstorming etc. |  |  |  |
| 6 M Method (If Applicable): - Man: Not Applicable. - Machine: Not Applicable - Measurement: Not Applicable - Material: Not Applicable - Method Not Applicable - Milieu (Environment) Not Applicable |  |  |  |
| Note: No any 6 M question shall be deleted in investigation, If any question is not applicable, then mention in ans. as “Not Applicable”. |  |  |  |
| 5 Why Approach (If Applicable): Q1. Why was the audit trail data not available on 28/10/2025? Ans: Because the system backup overwrote the existing file instead of archiving it separately. Q2. Why did the system overwrite the existing file? Ans: Because the file was not renamed before archiving, the system treated it as the same file and replaced it during the next backup cycle. Q3. Why was the file not renamed before archiving? Ans: Because IT did not follow the OEM vendor manual instructions during initial qualification and subsequent backup operations. Q4. Why did IT not follow the OEM vendor manual instructions? Ans: Because the instruction regarding mandatory file renaming before archiving was either overlooked, not documented properly, or not integrated into the IT SOPs/training. Q5. Why was the instruction overlooked or not integrated into SOPs? Ans: Because during initial qualification, vendor guidance was not fully captured in the qualification protocol and IT procedures, leading to a gap in compliance with OEM requirements. Note: Less or more than 5 questions may be asked based on the outcome of earlier questions. Conclusion: Based on the 5 why analysis, it is concluded that at initial qualification manual provided by the vendor regarding backup overwritten function point not included in archival procedure in IT SOP SOP/DP/IT/001. Hence inadvertently missed to backup file rename before next backup due. |  |  |  |
| Brainstorming: Not Applicable |  |  |  |
| Other Tool if Any: Not Applicable |  |  |  |
| Other Tool if Any: Not Applicable |  |  |  |
| Investigation Outcome: |  |  |  |
| During the review of the QA Audit trail of PH Conductivity Meter (I/QC/10) for month October 2025 it was identified that data is not available for Audit Trail PH Conductivity Meter of 28/10/2025. QA department reviews the data for direct impacting and indirect impacting system on scheduled basis. PH Conductivity Meter (I/QC/10) in QC department audit trail done on 17/11/2025. During audit trail it is observed that the audit trail data is missing for audit trail from date14/10/2025 to 28/10/2025. After that QA Personnel informed the QC and QA head and IT about these discrepancies. IT Personnel visited the system and checked last backup file available with server storage. Last backup was found of 27/10/2025 dated. As a corrective action IT Personnel immediately copy that backup file in backup PC. Again, QA reviewed the same date audit trail were discrepancy found and it is observed that data recovered for discrepancy dates however 28/10/2025 data for audit were not found. It is observed that IT Personnel taken schedule backup on dated 27/10/2025 as per SOP/DP/IT/001 and copied on server on dated 28/10/2025, but after that on dated 29/10/2025 PH Conductivity memory full error occurred. User department approach to IT and QA department for Archival data as per the defined procedure in SOP No. ………….. User department raise archived request, as per the request QA archived data on dated 29/10/2025, Before archived as per the procedure IT taken backup and keep at it is on dated 27/10/2025. IT next backup due on dated 02/11/2025 that time due to system overwritten functionality that file overwrite because of file was not renamed with last backup date on dated 29/10/2025, hence data missed to captured after last backup and before last archival which is 28/10/2025. It is concluded that, Data for 28/10/2025 was lost because of system overwrite functionality combined with IT Personnel’s failure to rename the backup file before archival. |  |  |  |
| Identified Root Cause/ Probable Cause: |  |  |  |
| Root cause I: Primary Root Cause Level 1: Human error Secondary Root Cause Level 2: Omission error Tertiary Root Cause Level 3: Responsibility boundaries ineffective. Root cause II: Primary Root Cause Level 1: Method / Procedure Secondary Root Cause Level 2: Document lacking details. Tertiary Root Cause Level 3: NA |  |  |  |
| Impact Assessment (System/ Document/ Product/ Equipment/Patient safety/Past batches): |  |  |  |
| System: The nonconformance is related to data missing for audit trail from PH Conductivity meter (I/QC/10) for date 28/10/2025 due to system over functionality.IT Personnel did not perform backup file rename activity however backup was successful before archived data. Data is available for the sample, batch only audit trail related data like user login, logout is missed for 28/10/2025 Document: As per the IT SOP No. SOP/DP/IT/001 data backup is taking on scheduled frequency basis. Every month 8/16/24 weekly backup (± 2 working days) taking by IT and for monthly backup on every month 2nd (± 5 working days) need to take backup. After taken backup IT Personnel fill the format SOP/DP/IT/F01-R01 Record for weekly and monthly backup. After that QA Personnel is verifying backup status and signing off the record log. For archival as when required need to execute the activity. For archival user department initiate the IT request. After that QA Personnel and IT Personnel execute the archival activity. Before archival IT Personnel taking backup and then QA Personnel free the memory space as per the requirement. Product: Archival is performing for the memory free space, due to memory full system not able to perform operation as per the requirement or data not able to log and system restrict to initiate next batch. Hence archival procedure follows as per requirement and regulations. There is no impact of data archival on products, however all data is available. Equipment: Selected Equipment for archival having some limitations for the memory as designed. However, whenever required to memory space in system and if it is restricted to initiate system operation then need to follow archival procedure for free memory from the system. Patient safety / Past Batches: Archival procedure execute for only free memory space from the system for the operation running smoothly and no impact is on compliance, and traceability, however all batches data is available with IT department, and it is retrieval in case of disaster management, and it is not impacted on patient safety. Hence there does not compromise patient safety due archival procedure. |  |  |  |
| Improve: Improve section covers the corrective actions Following checkpoint shall be considered as guidance only while finalizing the corrective actions, - Is a specific corrective Actions identified (including applicable Immediate Actions) to remediate the current issue or situation such that the associated system was returned to a state of control/compliance and also it will avoid the recurrence of deviation? - Is specific Corrective Actions identified for each root cause /substantiated probable root cause, as applicable? - Is the Corrective action assigned a unique number, responsible Personnel and due date so it can be tracked? Does the action describe what will be the expected outcome they can be verified? - Are the identified corrective actions achievable based on the information provided? |  |  |  |
| Corrective Action: |  |  |  |
| Upon identification of the non-conformance following corrective action shall be taken: - Identified the equipment / instrument for performing the data Archival applicability and list to be prepared. |  |  |  |
| Control: Control section covers the preventive actions Following checkpoint shall be considered while finalizing the preventive actions, - Were specific Preventive Actions identified for each root cause /substantiated probable root cause as applicable? - Was the Preventive Action linked the classification of the root cause and explanation given for how it will prevent recurrence? - Was the Preventive action assigned a unique number, responsible Personnel and due date so it can be tracked? Does the action describe and expected outcome can be verified? - Was Effectiveness Verification required or not and the rationale for either documented based on the quality impact of the incident after the investigation concluded? - Was an Interim Plan needed to ensure a state the control while the Preventive Actions were implemented? - Was rationale provided when no Preventive Action were identified? - Does the Final Comments section include rotational to support the conclusion of the investigation and CAPA. - Was each of the impact assessment fields completed correctly - Regulatory Impact, Regulatory notification Product Quality, Validation, Stability, Market/Clinical? - Does the Recommended Lot disposition match the conclusions of the investigation and Impact assessment? - Does the Conclusion include final decision and rationale other regulatory notification required? - CAPA required to release material or batches have been verified to be complete and closed prior to material or batch disposition. Any related CAPA that remain open, but have no impact on the material or batch release and a documented explanation is provided. - Does the Conclusion include a brief summary of root cause, final scope/impact including rationale for product/material disposition, impact assessment and relevant lot/material/equipment details? - Are the identified preventive actions achievable based on the information provided? |  |  |  |
| Preventive Action: |  |  |  |
| - SOP for Electronic Data Backup and Retrieval, SOP no. SOP/DP/IT/001 shall be revised for the procedural clarity for applicable data archival system that after every backup before every archival generated backup file shall be renamed with date to avoided overwrite data with the current data file. Interim Control: NA |  |  |  |
| Document Reviewed: - SOP for Electronic Data Backup and Retrieval, SOP no. SOP/DP/IT/001. - Deviation form: DEV/QC/25/002. - SOP/DP/QA/003-R04 - Vendor Manual |  |  |  |
| List of attachments (If applicable): |  |  |  |
| - Attachment No. I: Root cause categorization. - Attachment No. II: Vendor Service report - Attachment No. II: Audit trail review sheet of logbook |  |  |  |

| Prepared By (Sign/Date) |  |
| --- | --- |
| Reviewed By (Sign/Date) |  |
| Reviewed By (Sign/Date) |  |
| Reviewed By (Sign/Date) |  |
| Reviewed By (Sign/Date) |  |
| Reviewed By (Sign/Date) |  |
| Reviewed By (Sign/Date) |  |
| Reviewed By (Sign/Date) |  |
| Reviewed By (Sign/Date) |  |
| Reviewed By (Sign/Date) |  |
| Approved By QA (Sign/Date) |  |