| Date: | 29/09/2025 | Deviation No. | DEV/QC/25/002 |
| --- | --- | --- | --- |
| Investigation tool used: 6M 5 Why Brainstorming |  |  |  |
| Other Tools (If any): |  |  |  |
| Not Applicable |  |  |  |
| Details Investigation: |  |  |  |
| Define: |  |  |  |
| The following checks shall be considered while writing the “Define” section. - Clearly define what happens actually. - Explain what is different than expected. - Mention the location at where the deviation has occurred. - Date/time of deviation occurrence and date/time of detection. - Mention the name of personnel who is involved in the deviation. - Mention initial scope of deviation (i.e., impacted product/Material/Equipment/System//Batches/etc.) While processing the chromatographic data of sequence 021_0100.lcb on 23/09/2025 at 17:17 hrs. which was initiated for calibration parameter i.e. “Wavelength Accuracy of Detector” of HPLC instrument I/QC/021 at Instrument Lab–I (FF-17), the analyst (Emp. ID: 952) observed that the minimum absorbance area was obtained at 242 nm, which is outside the defined acceptance criteria of the SOP (245 nm ± 2 nm; i.e., 243 nm to 247 nm). As per SOP: SOP/DP/QC/059 Revision No.: R04, Title: Operation, Calibration, Maintenance and Data Management of High-Performance Liquid Chromatography (HPLC) with Lab Solutions Software, (Section No.7.16.16.5) Acceptance Criteria: The maxima area at 205 nm ± 2 and 273 nm ± 2, the minimum area at 245 nm ± 2 |  |  |  |

Details Investigation:

Measure: The following checks shall be considered while writing the “Measure” section. - Does the summary provide relevant facts and data/information that was reviewed including: environment, process/product history, personnel info (title and job title), controls/control limits, etc. - Is a summary of the analysis of the factors and data provided? - Is a conclusion statement of the analysis and review provided? - If there were Regulatory Notification, were details provided? - Is the report written in a logical flow and easily understood by the reader? Calibration of HPLC instrument I/QC/021 scheduled on 04/10/2025 and as per SOP No.: SOP/DP/QC/012, Revision No: R01, Title: Procedure for Calibration and Preventive Maintenance of QC Instruments and as per SOP, six monthly calibrations can be initiated before 15 working days from the scheduled date and to be closed on or before the scheduled date. Since, calibration of HPLC instrument I/QC/021 was due on 04/10/2025, hence calibration of the HPLC system planned and work allotted to analyst Emp. Id. 952 on 21/09/2025. As per plan, analyst generated the analytical reference number (A.R. No NR/25/0092) and verified all the pre requisites before initiation of calibration and after collecting all the requisites initiated the physical parameter for calibration of HPLC instrument I/QC/021 on dated 21/09/2025. As per SOP: SOP/DP/QC/059 Revision No.: R04, Title : Operation, Calibration, Maintenance and Data Management of High Performance Liquid Chromatography (HPLC) with Lab Solutions Software, analyst (Emp. Id: 962) initially performed the ‘Flow Rate Accuracy’ test for the set flow 0.2 mL/min, 0.5 mL/min, 1.0 mL/min, 2.0 mL/min, 3.0 mL/min and obtained results are found well within the acceptance criteria. (Refer Attachment No. I). Further, performed the temperature accuracy test of ‘Auto Sampler compartment’, and ‘Column Oven Compartment’ of systems on 21/09/2025 and obtained results that were found well within the acceptance criteria, hence analyst proceeded with the remaining calibration parameters (Refer Attachment No. I). On 22/09/2025, analyst (Employee ID :952) initiated the ‘Gradient Composition’ test of HPLC instrument I/QC/021 and after readiness of all the prerequisite, analysts prepared the sample sequence ‘021_0099’ and submitted to QC reviewer for review. QC Reviewer reviewed the sample set sequence as per HPLC readiness checklist (Format No.: SOP/DP/QC/059/F01-R00) and given clearance to analyst for initiation of sequence and further, analyst initiated the sequence ‘021_0099’ on 22/09/2025 for ‘Gradient Composition’. After successful completion of the initiated sequence for ‘Gradient Composition’ (Sequence No. 021_0099; Data File Nos. 021_002482 to 021_002487) on 22/09/2025 at 17:42 hours, the analyst subsequently prepared the mobile phase for the next calibration parameter i.e. ‘Wavelength Accuracy of Detector’. On 22/09/2025 at 18:17:49 hours, the analyst (Employee ID: 952) created Sequence No. 021_0100 (Data File Nos. 021_002488 to 021_002529) for the ‘Wavelength Accuracy’ parameter on HPLC Instrument I/QC/021 (Refer Attachment-II) and submitted the sequence to the QC Reviewer for review.” QC reviewer verified the mobile phase preparation, caffeine standard preparation and column details as per the details given below, • Mobile Phase Preparation: Mixed 450 mL of Acetonitrile with 2550 mL of Water. • Column Used: HPLC/COL/001-00.” Preparation of Standard Solution: Weighed 99.83 mg of caffein standard into 100 ml volumetric flask, dissolved and diluted to volume with mobile phase (1000 ppm). Preparation of Standard Solution: (50 ppm): Diluted 5 ml of 1000 ppm solution to 100 ml and diluted to volume with mobile phase (50 ppm). Analyst prepared the mobile phase and caffeine standard considering the subsequent calibration sequences, for which the mobile phase, column, and standard preparation remained the same. The subsequent calibration sequences include ‘Linearity of Auto Injector’, ‘Linearity of Detector Response’, and ‘Volume Precision and Carry Over’ QC Reviewer reviewed the sample set sequence as per HPLC readiness checklist (Format No.: SOP/DP/QC/059/F01-R00) and given clearance to analyst for initiation of sequence and further, analyst initiated the sequence ‘021_0100’ on 22/09/2025 for ‘Wavelength Accuracy’. Subsequently, the analyst (Employee ID: 952) created the below sequences for further calibration parameter and submitted to QC reviewer for review, Sr. No.; Calibration Parameter; Batch File No.; Data File No.; Batch File Creation time. 1.; Linearity of Auto Injector; 021_0101; 021_002530 to 021_002559; 22/09/25,18:41:07. 2.; Linearity of Detector Response; 021_0102; 021_002560 to 021_002571; 22/09/25,18:58:23. 3.; Volume Precision and Carry Over; 021_0103; 021_002572 to 021_002581; 22/09/25,19:14:12. QC Reviewer reviewed the above sequences as per HPLC readiness checklist (Format No.: SOP/DP/QC/059/F01-R00) and given clearance to analyst for initiation of sequences. The analyst registered the following batch files in the queue one by one for execution. The details are as follows: (Refer attached Audit Trail), Sr. No.; Calibration Parameter; Batch File No.; Data File No.; Batch File Registration in queue (date & time ). 1.; Linearity of Auto Injector; 021_0101; 021_002530 to 021_002559; 22/09/25,19:46:41. 2.; Linearity of Detector Response; 021_0102; 021_002560 to 021_002571; 22/09/25,19:47:05. 3.; Volume Precision and Carry Over; 021_0103; 021_002572 to 021_002581; 22/09/25,19:47:21. On 23/09/2025 at 17:17 hrs, while processing the chromatographic data of sequence 021_0100.lcb which was initiated for calibration parameter “Wavelength Accuracy of Detector” of HPLC instrument I/QC/021 at Instrument Lab–I (FF-17), the analyst (Emp. ID: 952) observed that the minimum absorbance area was obtained at 242 nm, which is outside the defined acceptance criteria of the SOP (245 nm ± 2 nm; i.e., 243 nm to 247 nm) hence to investigate occurred nonconformance initiated the deviation. Since the observed wavelength (242 nm) is out of the defined acceptance range, the following immediate corrective action were taken: - Mobile phase supply lines were checked for trapped air bubbles, and it was noted that no physical air bubble was observed in the mobile phase supply line. - The column compartment was verified for any leakage, and it was noted that no leakage was observed. - The retention time of the caffeine peak was verified and observed at Caffeine peak eluted approximately at 6 minutes, whereas it is expected to be around 3 minutes. - A preliminary review of previous calibration file (021_0099.lcb) which was performed for Gradient Gradient Composition’ test, which confirmed that results were found within acceptance criteria. - Verified the pressure graph of all the completed injection and noted that pressure was stable up to the run of Data file 021_002488 to 021_002503, however it was observed that a pressure drop occurred in Data File Nos. 021_002504 to 021_002529, which were initiated for the wavelength accuracy parameter, which led to the shifting of retention time of caffeine was observed at approximately 6.9 min, whereas it was expected to be approximately 3.4 min. For details refer Attachment -IV. - The details retention time of caffein as below, Sr. No.; Data File No.; Retention Time. 021_002489; 3.492. 021_002491; 3.492. 021_002493; 3.492. 021_002495; 3.492. 021_002497; 3.492. 021_002499; 3.492. 021_002501; 3.492. 021_002503; 3.492. 021_002505; 6.958. 021_002507; 6.958. 021_002509; 6.958. 021_002511; 6.958. 021_002513; 6.958. 021_002515; 6.967. 021_002517; 6.958. 021_002519; 6.967. 021_002521; 6.967. 021_002523; 6.967. 021_002525; 6.967. - Based on the occurred non-conformance, verified the status of sequences which were loaded in the queue i.e. 021_0101, 021_0102 and 021_0103 and noted that sequence 021_0101 was under running, however remaining sequences i.e. 021_0102 and 021_0103 were in queue. - Further verified the pressure of completed injections of batch file 021_0101 (Data File Nos. 021_002530 to 021_002559), which was initiated for the linearity of the auto injector calibration parameter, and noted that a pressure drop occurred in all data files from 021_002530 to 021_002559. The retention time of the caffeine peak was observed at approximately 6 minutes, whereas the expected retention time was around 3 minutes, therefore as an immediate action by considering the cascading impact on the subsequent batch files 021_0102 and 021_0103, which were in the queue for analysis and initiated for accuracy of the auto-injector and volume precision, have been removed from the queue. Since an exact root cause was not identified for the occurred nonconformance during the preliminary investigation, hence the observed nonconformance was escalated to the HPLC service engineer for further evaluation as the service engineer was available on site on 23/09/2025 for planned preventive maintenance activity of HPLC (I/QC/022). Based on the reported discrepancy, service engineer verified the HPLC system I/QC/021, and during verification, the following observations were made (Refer Service Engineer Service Report): - The service engineer checked the sequence pressure and found that initially, the injections were stable and the retention time (RT) was consistent. However, at a particular point during the injections (Data File Nos. 021_002504), a drop in pressure was observed, which impacted the retention time. - Further, purged the HPLC system I/QC/021 properly and checked the solvent filters for any blockage, and no issues were found. - The check valves were inspected and found to be working satisfactorily, and the pressure was monitored and observed to be stable. - No other physical or mechanical abnormalities were detected in the HPLC system. - Based on the reviewed data, the most probable cause of the pressure drop is an air bubble trapped in the pump, specifically stuck in the check-inlet tubing, which impacted the system pressure drop which led to occurrence inaccurate area response of wavelength accuracy standard solution. Based on the occurred nonconformance, exact root cause could not be identified for the occurred nonconformance, however the most probable cause for could be ‘Instrument error’ i.e. Air bubble trapped in the pump, specifically stuck in the check-inlet tubing, which impacted the system pressure drop which led to occurrence of inaccurate area response of wavelength accuracy standard solution. Therefore, based on the most probable cause it was recommended to purge the HPLC system I/QC/021, properly and initiate the wavelength accuracy of detector test after getting the stabilized pressure with Sequence No. 021_0102. (Note: Sequence 021_0102 and 021_0103, were initially created for ‘Accuracy of the auto-injector’ and ‘Volume precision’, however due to occurred nonconformance both the sequences not proceeded, hence to follow the sequential data file numbering, sequence 021_0102 was modified for the ‘Wavelength accuracy of detector’ and new sequence i.e. 021_0104 created for ‘Accuracy of the auto-injector’. As per recommended action, wavelength accuracy of detector test of the HPLC instrument i.e. I/QC/021 was planned and work was allotted to analyst Emp. Id. 952 on 24/09/2025. As per the recommended action, the analyst purged the system and verified the system pressure before initiation of the ‘Wavelength Accuracy’ test. As per the plan, the analyst verified all the prerequisites before initiation of the ‘Wavelength Accuracy’ test and after collecting all the requisites initiated the preparation for the ‘Wavelength Accuracy’ test of HPLC instrument I/QC/021. After ensuring the stabilized pressure, analyst (Employee ID: 952) modified the sequence No. 021_0102 (Data File Nos. 021_002560 to 021_002601) for the ‘Wavelength Accuracy’ parameter on HPLC Instrument I/QC/021 (Refer Attachment-VI) and submitted the sequence to the QC Reviewer for review.” QC reviewer verified the mobile phase preparation, caffeine standard preparation and column details as per the details given below, • Mobile Phase Preparation: Mixed 450 mL of Acetonitrile with 2550 mL of Water. • Column Used: HPLC/COL/001-00.” Preparation of Standard Solution: Weighed 100.03 mg of caffein standard into 100 ml volumetric flask, dissolved and diluted to volume with mobile phase (1000 ppm). Preparation of Standard Solution: (50 ppm): Diluted 5 ml of 1000 ppm solution to 100 ml and diluted to volume with mobile phase (50 ppm). Analyst prepared the mobile phase and caffeine standard considering the subsequent calibration sequences, for which the mobile phase, column, and standard preparation remained the same. The subsequent calibration sequences include ‘Linearity of Auto Injector’, ‘Linearity of Detector Response’, and ‘Volume Precision and Carry Over’ QC Reviewer reviewed the sample set sequence as per HPLC readiness checklist (Format No.: SOP/DP/QC/059/F01-R00) and given clearance to analyst for initiation of sequence and further, analyst initiated the sequence ‘021_0102’ on 24/09/2025 for ‘Wavelength Accuracy’. After completion of sequence analyst verified the pressure for all the completed injection and noted that pressure was stable throughout the sequence and further analyst processed the data on 24/09/2025. Chromatographic data of sequence reviewed and reported on 25/09/2025 and while reporting the results for ‘Wavelength Accuracy’ test analyst observed that all the obtained results are found well within the acceptance criteria. For details refer Attachment – I. Based on the outcome of investigation and obtained results it is revealed that, in the initial analysis, air bubble was trapped in the pump, specifically stuck in the check-inlet tubing, which was impacted the system pressure drop which led to occurrence inaccurate area response of wavelength accuracy standard solution and consequently minimum absorbance area was obtained at 242 nm, which is outside the defined acceptance criteria of the SOP (245 nm ± 2 nm; i.e., 243 nm to 247 nm).

Analyze: Various investigation tools used to analyze the root cause like 6M, 5 Why, Personnel interview, Brainstorming etc.

6 M Method (If Applicable): - Man: Q1: Verify if any new personnel (officer or operator) involved during manufacturing process? Ans. Not Applicable Q 2. Is he/she qualified enough to perform assigned operation? Ans. Not Applicable Q 3. Is he/she experienced enough to perform assigned operation? Ans. Not Applicable Q 4. Was there proper training given to perform the assigned task? Ans. Not Applicable Q 5. Were procedures available to perform assigned task? Ans. Not Applicable Q 6. Whether individual adhered to respective procedure? Ans. Not Applicable Q 7. Did the environment influence the actions of the individual? Ans. Not Applicable Q 8. Was the procedure / specifications properly interpreted? Ans. Not Applicable Q 9. Was too much judgment required to perform the task? Ans. Not Applicable Q 10. Are there interruptions in the workplace, which interferes in the functioning of the individual? Ans. Not Applicable Q 11. Are sufficient numbers of people available to carry out the activity? Ans. Not Applicable Q 12. Is there any time pressure? Ans. Not Applicable Q 13. Verify if there is any evidence that any change in operator during manufacturing process. Ans. Not Applicable *Q14: Not Applicable Ans. Not Applicable *Any other question related to the particular. Conclusion: Not Applicable

- Machine: Q 1. Whether machine used for activities are qualified? Ans. Not Applicable Q 2. Were correct tools/ change parts used? Ans. Not Applicable Q 3. Does it meet process capabilities and production requirements? Ans. Not Applicable Q 4. Is the equipment affected by the environment or any other factor? Ans. Not Applicable Q 5. Is the equipment being properly maintained? Whether PM/ Calibration are carried out as per specified frequency? Ans. Not Applicable Q 6. Was the machine properly programmed? Ans. Not Applicable Q 7. Is it clearly labelled? Ans. Not Applicable Q 8 Whether machine parameters are set as per instructions provided? Ans. Not Applicable Q 9. Is there any breakdown during batch processing? Ans. Not Applicable Q 10. Is equipment used within operating range? Ans. Not Applicable Q 11. Was there any discrepancy during processing / operation? Ans. Not Applicable Q 12. Verify the SOPs are available for operation of all equipment’s / instruments? Ans. Not Applicable *Q 13. Not Applicable Ans. Not Applicable *Any other question related to the particular. Conclusion: Not Applicable

- Measurement: Q 1. Whether adequate in-process checks are in place for proper functioning? Ans. Not Applicable Q 2. Were the proper gauge/instruments used to measure the part, process, chemical, compound, etc.? Ans. Not Applicable Q 3. Were any discrepancies or failures observed during batch processing? Ans. Not Applicable Q 4. Does measuring devices have a valid calibration date? Ans. Not Applicable Q 5. Is there any SOP available for calibration? Ans. Not Applicable Q 6. Was the instrument calibrated in required range? Ans. Not Applicable Q 7. Is the calibration done as per specified frequency? Ans. Not Applicable Q 8. Whether personal variation affects performance of measuring instruments? Ans. Not Applicable Q 9. Did the environment or any other factors influence the measurements taken? Ans. Not Applicable Q 10. Was the instrument / equipment validated? Ans. Not Applicable *Q 11. Not Applicable Ans. Not Applicable *Any other question related to the particular Conclusion: Not Applicable

- Material: Q 1. Whether material purchased from approved supplier/vendor? Ans. Not Applicable Q 2. Is there a problem associated with the material? Ans. Not Applicable Q 3. Was the material approved by QC / QA? Ans. Not Applicable Q 4. Is all needed information available and accurate? Ans. Not Applicable Q 5. Is a MSDS available? Ans. Not Applicable Q 6. Is the supplier’s process defined and controlled? Ans. Not Applicable Q 7. Was the material handled properly during receipt, transit, storage & dispensing? Ans. Not Applicable Q 8. Is there any change in supplier or supplier’s site? Ans. Not Applicable Q 9 Is there any change in quality parameter as compared to previous consignment? Ans. Not Applicable Q 10. Is sampling procedure followed as per SOP? Ans. Not Applicable Q 11. Is there any abnormal remark in receipt checklist? Ans. Not Applicable Q 12. Is material released under Deviation/OOS/OOT? Ans. Not Applicable Q 13. Is vendor/supplies testing result available? Ans. Not Applicable *Q 14. Not Applicable Ans. Not Applicable *Any other question related to the particular. Conclusion: Not Applicable

- Method: Q 1. Is approved document/ procedure available for processing? Ans. Not Applicable Q.2. Was the process or method changed? Ans. Not Applicable Q 3. Are activities being performed as per approved procedure mentioned in batch doc/ SOP / etc.? Ans. Not Applicable Q 4. Is there any deviation with respect to process or method? Ans. Not Applicable Q 5. Were the individual trained to perform allocated activity? Ans. Not Applicable Q 6. Are the instructions to be followed during execution of activity clearly written? Ans. Not Applicable Q 7. Was the proper labeling in place? Ans. Not Applicable Q 8. Is handling/packaging adequately specified? Ans. Not Applicable Q 9. Are the lighting / ventilation adequate? Ans. Not Applicable Q 10. Was a risk assessment ever performed? Any outcomes from this? Ans. Not Applicable Q 11. Any other questions? Ans. Not Applicable Q 12. Is in process testing done as per defined frequency? Ans. Not Applicable Q 13. Is line clearance followed adequately? Ans. Not Applicable Q 14. Whether applicable validations are in place? Ans. Not Applicable *Q 15. Not Applicable Ans. Not Applicable *Any other question related to the particular. Conclusion: Not Applicable

- Milieu (Environment): Q 1. Is the process under consideration affected by humidity, vibration, noise, lighting, etc.? Ans. Not Applicable Q 2. Were the environmental conditions like temperature, humidity etc. within limits? Ans. Not Applicable Q 3. Is the process under consideration affected by environment (e.g. RH sensitive / light sensitive)? Ans. Not Applicable Q 4. Are there interruptions in the workplace which interferes in the functioning of the individual? Ans. Not Applicable Q 5. Is the area qualified and re-qualified as per frequency? Ans. Not Applicable Q 6. Is AHU qualified and is as per production requirement? Ans. Not Applicable Q 7. Is there any evidence of power failure during the manufacturing process? Ans. Not Applicable *Q 8. Not Applicable Ans. Not Applicable *Any other question related to the particular. Conclusion Not Applicable

Note: No any 6 M question shall be deleted in the investigation, if any question is not applicable, then mention in ans. as “Not Applicable”.

5 Why Approach (If Applicable): 1. WHY: ‘Wavelength Accuracy’ test was found out of calibration (OOC)? Ans. ‘Wavelength Accuracy’ test found out of calibration, because minimum absorbance obtained at 242 nm, which is outside the defined acceptance criteria of the SOP (245 nm ± 2 nm; i.e., 243 nm to 247 nm). 2. WHY: Minimum absorbance area was obtained at 242 nm? Ans. Due to drop in the HPLC system pressure at Data File Nos. 021_002504 to 021_002529, inaccurate area response occurred for the wavelength accuracy standard solution consequently minimum absorbance area was obtained at 242 nm 3. WHY: Why HPLC system pressure dropped? Ans: Probably, air bubble was trapped in the pump, specifically stuck in the check-inlet tubing, which led to drop of HPLC system pressure. 4. WHY: Not Applicable? Ans. Not Applicable *5. WHY: Not Applicable Ans. Not Applicable *Note: Less or more than 5 questions may be asked based on the outcome of earlier questions. Conclusion: Based on the 5 Why Approach outcome, it is concluded that nonconformance was occurred due to ‘Instrument error’ i.e. Air bubble trapped in the pump, specifically stuck in the check-inlet tubing, which impacted the system pressure drop which led to occurrence of inaccurate area response of wavelength accuracy standard solution.

| Investigated Person Details: |  |  |  |
| --- | --- | --- | --- |
| Name | Dhananjay Ibhad | Employee ID | 952 |
| Department | Quality Control | Date | 29/09/2025 |
| Designation | Executive |  |  |

| Interviewer Name | Nilesh Dushing | Interviewer Employee ID | 616 |
| --- | --- | --- | --- |
| Designation | Dy. Manager |  |  |

| Interviewer Name | Pravin Kolkand | Interviewer Employee ID | 622 |
| --- | --- | --- | --- |
| Designation | Sr. Executive. |  |  |

Personal Interview (If Applicable): Q1. Not Applicable Ans. Not Applicable Q2. Not Applicable Ans. Not Applicable Q3. Not Applicable Ans. Not Applicable Q4. Not Applicable Ans. Not Applicable Q5. Not Applicable Ans. Not Applicable Q6. Not Applicable Ans. Not Applicable Q7. Not Applicable Ans. Not Applicable *Q8. Not Applicable *Note: Number of questions shall be decided based on the given answers and type of the deviation.

Signatures:

| Interview Person Details: |  |  |  |
| --- | --- | --- | --- |
| Name | NA | Employee ID | NA |
| Department | NA | Date of Interview | NA |
| Designation | NA | Sign/Date | NA |

| Interviewer Name | NA | Interviewer Employee ID | NA |
| --- | --- | --- | --- |
| Designation | NA | Sign/Date | NA |

| Interviewer Name | NA | Interviewer Employee ID | NA |
| --- | --- | --- | --- |
| Designation | NA | Sign/Date | NA |

Brainstorming: Not Applicable

Other Tool if Any: Not Applicable

Other Tool if Any: Not Applicable

Investigation Outcome: The investigation has been carried out by using the ‘5 Why’ investigation tool. The investigation was driven through the DMAIC methodology. Based on the initial risk assessment, the deviation has been categorized the deviation as ‘Minor’. The nonconformance related to the occurrence of minimum absorbance obtained at 242 nm, which is outside the defined acceptance criteria of the SOP (245 nm ± 2 nm; i.e., 243 nm to 247 nm), therefore detailed investigation has been carried out, and based on the investigation outcome root cause is identified as ‘Instrument error’ i.e. Air bubble trapped in the pump, specifically stuck in the check-inlet tubing, which impacted the system pressure drop which led to occurrence of inaccurate area response of wavelength accuracy standard solution. Based on the identified root cause, performed the impact assessment found that there is no impact on the sample analyzed on HPLC instrument I/QC/021, as the results obtained for analyzed samples were found well within the acceptance criteria and there is no any similar type of error occurred while performing the sample analysis, hence there is no impact on the previously performed analysis. Also, before occurrence of nonconformance in the wavelength accuracy test, other calibration parameter was performed i.e. ‘Gradient Composition’, Flow rate accuracy test and temperature accuracy and obtained results are found well within the acceptance criteria, hence there is no impact on the performance of instrument, as the HPLC instruments I/QC/021 was utilized in the various analysis, however, there were no any issue reported related to ‘Wavelength’ issue, hence no impact on the performance of HPLC instruments I/QC/021. Based on the nature of nonconformance, chronology of events, and details investigation, further root cause categorization has been done as per format for “Root Cause Categorization” SOP/DP/QA/008-F03. (Refer Attachment No. VII)

Identified Root Cause/ Probable Cause:

- Based on the detailed investigation, below mention probable root cause were identified Primary Root Cause Level 1: Instrument error Secondary Root Cause Level 2: Mechanical Failure Third Root Cause Level 3: NA

Impact Assessment (System/ Document/ Product/ Equipment/Patient safety/Past batches):

System: The occurred non-conformance is pertaining to the occurrence of minimum absorbance obtained at 242 nm, which is outside the defined acceptance criteria of the SOP (245 nm ± 2 nm; i.e., 243 nm to 247 nm) due to ‘Instrument error’ i.e. Air bubble trapped in the pump, specifically stuck in the check-inlet tubing, which impacted the system pressure drop which led to occurrence of inaccurate area response of wavelength accuracy standard solution, therefore impact on ‘System’ is not anticipated. Document: While performing the “Wavelength Accuracy of Detector” analyst followed the correct practice as per SOP No.: SOP/DP/QC/059, Title: Operation, Calibration, Maintenance and Data Management of High-Performance Liquid Chromatography (HPLC) with Lab Solutions Software’, however non-conformance occurred due to ‘Instrument error’ i.e. Air bubble trapped in the pump, hence there is no impact on the ‘Document’. Product: Based on the reported nonconformance, performed the impact assessment and noted that there is no any similar type of error occurred while performing the sample analysis and results obtained for analyzed samples were found well within the acceptance criteria, hence there is no impact on the previously performed analysis and quality of ‘Product’. Instrument: The occurred non-conformance is pertaining to the occurrence of minimum absorbance obtained at 242 nm, which is outside the defined acceptance criteria of the SOP (245 nm ± 2 nm; i.e., 243 nm to 247 nm) due to ‘Instrument error’ i.e. Air bubble trapped in the pump before and it may happen due to mechanical failure, however it does not have impact on the performance of instrument, as before occurrence of nonconformance in the wavelength accuracy test, other calibration parameter was performed i.e. ‘Gradient Composition’, Flow rate accuracy test and temperature accuracy and obtained results are found well within the acceptance criteria. Patient safety/Past Batches: The occurred non-conformance is pertaining to the occurrence of minimum absorbance obtained at 242 nm while performing the calibration of HPLC instrument, hence based on the reported nonconformance, performed the impact assessment and noted that there is no any similar type of error occurred while performing the sample analysis and results obtained for analyzed samples were found well within the acceptance criteria, hence there is no impact on the previously performed analysis and quality of ‘Product’, Patient safety/Past Batches.

Improve: Improve section covers the corrective actions Following checkpoint shall be considered as guidance only while finalizing the corrective actions, - Is a specific corrective Actions identified (including applicable Immediate Actions) to remediate the current issue or situation such that the associated system was returned to a state of control/compliance and also it will avoid the recurrence of deviation? - Is specific Corrective Actions identified for each root cause /substantiated probable root cause, as applicable? - Is the Corrective action assigned a unique number, responsible person and due date so it can be tracked? Does the action describe what will be the expected outcome they can be verified? - Are the identified corrective actions achievable based on the information provided?

Corrective Action:

The investigation has been driven through the DMAIC methodology. During the investigation ‘5 Why’ tool used to find out the root cause and based on the detailed investigation, it has been concluded that the nonconformance was triggered due to ‘Instrument error’ i.e. Air bubble trapped in the pump, specifically stuck in the check-inlet tubing, which impacted the system pressure drop which led to occurrence of inaccurate area response of wavelength accuracy standard solution. Hence, as a corrective action “Wavelength Accuracy of Detector” test of the HPLC system I/QC/021 has been repeated by purging the HPLC system and verified the system pressure before initiation of the test with Sequence No. 021_0102 and as an outcome obtained results are found well within the acceptance criteria. For details refer Attachment – I, hence chromatographic data generated from Sequence No. 021_0102 (Data File Nos. 021_002560 to 021_002601) considered for reporting and Chromatographic data generated from sequence ‘021_0100’ (Data file 021_002488 to 021_002529) and sequence 021_0101 (Data File Nos. 021_002530 to 021_002559) has been invalidated. All necessary corrective actions have been initiated and completed; hence no further corrective actions are anticipated.

Control: Control section covers the preventive actions Following checkpoint shall be considered as guidance only , - Is specific Preventive Actions identified for each root cause /substantiated probable root cause as applicable? - Is the Preventive Action linked the classification of the root cause and explanation given for how it will prevent occurrence? - Is the Preventive action assigned a unique number, responsible person and due date so it can be tracked? Does the action describe and expected outcome can be verified? - Is an Interim Plan needed to ensure a state the control while the Preventive Actions were implemented? - Is rationale provided when no Preventive Action were identified? - Does the Final Comments section include rotational to support the conclusion of the investigation and CAPA - Is each of the impact assessment fields completed correctly — Regulatory Impact, Regulatory notification Product Quality, Validation, Stability, Market/Clinical? - Does the Recommended Lot disposition match the conclusions of the investigation and Impact assessment? - Does the Conclusion include final decision and rationale other regulatory notification required? - CAPA required to release material or batches have been verified to be complete and closed prior to material or batch disposition. Any related CAPA that - remain open, but have no impact on the material or batch release and a documented explanation is provided. - Does the Conclusion include a brief summary of root cause, final scope/impact including rationale for product/material disposition, impact assessment and relevant lot/material/equipment details? - Are the identified preventive actions achievable based on the information provided?

Preventive Action:

As an outcome of investigation, it was revealed that nonconformance was occurred due to ‘Instrument error’ i.e. Air bubble trapped in the pump, specifically stuck in the check-inlet tubing, which impacted the system pressure drop which led to occurrence of inaccurate area response of wavelength accuracy standard solution, hence by considering the isolated instance preventive action not anticipated for the ocuured nonconformance, however to avoide the reoccurence, awarness training has been imparted to all the concerned personnel and instructed them to purge the HPLC system properly, verify the system pressure before initiation of analysis and initiate the sequence only after getting the stabilized system pressure.

Document Reviewed: - SOP No.: SOP/DP/QC/059 - Calibration Record: I/QC/021

List of attachment (If applicable):

Attachment – I : Calibration Results. Attachment – II : Sequence - Wavelength Accuracy. Attachment – III : Audit Trail Attachment - IV : Pressure graph & Chromatographic Data. Attachment – V : Service Report Attachment – VI : Sequence - Wavelength Accuracy (Updated). Attachment – VII : Root Cause Categorization

| Prepared By (Sign/Date) | Reviewed By (Sign/Date) | Reviewed By (Sign/Date) | Approved By QA (Sign/Date) |
| --- | --- | --- | --- |