| Date: | 09/04/2026 | Deviation No. | DEV/PK/25/002 |
| --- | --- | --- | --- |
| Investigation tool used: 6M 5 Why Brainstorming |  |  |  |
| Other Tools (If any): |  |  |  |
| Not applicable |  |  |  |
| Define: |  |  |  |
| Following checks shall be considered while writing the “Define” section. - Clearly define what happens actually. - Explain what is different than expected. - Mention the location where the deviation has occurred. - Date/time of deviation occurrence and date/time of detection. - Mention the name of personnel who is involved in the deviation. - Mention initial scope of deviation (i.e., impacted product/Material/Equipment/System/Batches/etc.) On 04/03/2026 at approximately 16:20 hrs., while reviewing the temperature data logs (from 06/02/2026 00:03:47 to 04/03/2026 09:23:36) of the Intermediate Walk-in Cold Room (2°C to 8°C), Equipment ID: E/PK/010, located in the packing area (GF-122), it was identified by a packing colleague (Employee ID: 598) that temperature data from 06/02/2026, 03:47 to 06/02/2026, 15:31 was not available in the data log report. This non-conformance was identified at the packing office (GF-109). As per site SOP No. SOP/DP/PK/013 (Operation and Cleaning of Intermediate Walk-in Cold Room), point 7.4.2 states that “The intermediate cold room temperature will be recorded in the system through Kyoshi software.” However, temperature data for the period from 06/02/2026, 03:47 to 06/02/2026, 15:31 was not recorded, which is a deviation from the SOP. The non-conformance occurred in the packing area where the Intermediate Walk-in Cold Room (E/PK/010) is located. At the time of the non-conformance, semi-finished and finished product batches, as well as visual inspection defect kits, were stored in the Intermediate Walk-in Cold Room. Although the non-conformance was identified, the Intermediate Walk-in Cold Room temperature was found to be within acceptable limits when verified at the last available instance, and no temperature excursion alarms occurred during the period when data was not logged in the system (Kyoshi application). Additionally, the daily MKT value on 06/02/2026 was 5.35 °C, which is within the acceptable range of 2 °C to 8 °C. Hence, there is no impact on the stored product/materials. and regulatory notification is not required |  |  |  |
| Details Investigation: |  |  |  |
| Measure: Following checks shall be considered while writing the “Measure” section. - Does the summary provide relevant facts and data/information that was reviewed including: environment, process/product history, personnel info (title and job title), controls limits, etc. - Is a summary of the analysis of the factors and data provided? - Is a conclusion statement of the analysis and review provided? - If there were Regulatory Notification, were details provided? - Is the report written in a logical flow and easily understood by the reader? The Intermediate Walk-in Cold Room (Equipment ID: E/PK/010) is installed in the packing area (GF-122) for the storage of semi-finished goods (temperature range: 2°C to 8°C). SOP No. SOP/DP/PK/013 (Operation and Cleaning of Intermediate Walk-in Cold Room) is in place for the cleaning and operation of the Cold Room. As per this SOP, the intermediate cold room temperature data shall be recorded in the Kyoshi application. However, on 06/02/2026, temperature data was not recorded in the system from 03:47 to 15:31. On 06/02/2026, at approximately 10:30 hrs. it was identified by a packing colleague (Employee ID: 1174) that the Cold Room HMI time was mismatched with the current time. At approximately 15:30 hrs. on the same day, the HMI display showed the time and date as 05/02/2026, 18:30 hrs. (approximate). Consequently, a work order (WO/PK/26-005, refer attachment no. I) was initiated to align the HMI time with the current time. The HMI time was updated by an IT colleague (Employee ID: TR0189) at approximately 15:40 hrs. on the same day. Thereafter, temperature data recording in the Kyoshi application was resumed. The work order was subsequently received by the packing colleague (Employee ID: 598) on 06/02/2026 to complete the “evaluation and verification of user” section in the work order form. As part of user verification, the HMI time and temperature data logging were monitored continuously until 04/03/2026. During this monitoring period, the HMI time remained accurate with no mismatches observed. Upon reviewing the temperature, alarm, and audit trail reports on 04/03/2026, it was confirmed that temperature data was properly logged in the Kyoshi application after the HMI time update (from 06/02/2026, 15:32 hrs. to 04/03/2026, 09:23 hrs.). However, temperature data during the period of HMI time mismatch, i.e., from 06/02/2026, 03:47 to 06/02/2026, 15:31, was not recorded in the Kyoshi application or captured in the reports. At the time of the non-conformance, semi-finished and finished product batches, as well as visual inspection defect kits, were stored in the Intermediate Walk-in Cold Room. Additionally, it was identified that a communication failure alarm occurred 7 times continuously on 06/02/2026 between 02:08:50 hrs. and 14:10:02 hrs. Following the first occurrence of this communication failure, temperature data logging in the Kyoshi application may have been halted. There is not any non-conformance identified in audit trail report (refer attachment no. II). The identified non-conformance was communicated to the respective User Head and QA Head, and a deviation was logged on 05/03/2026 at 13:05 hrs. to conduct a detailed investigation and identify the root cause. The non-conformance was also communicated to the Original Equipment Manufacturer (OEM), M/s Guru Control System, to support the investigation and assist in determining the actual root cause. Based on the initial risk assessment, the deviation was categorized as “Major”; therefore, a preliminary investigation was carried out to identify the most probable root cause. Refer to Attachment No. III for preliminary investigation details, including the most probable root cause. The Blue Star make walk-in cold rooms and deep freezers located in the warehouse and QC areas were also checked for any similar type of non-conformance, and no discrepancy was found. As part of the investigation, the cold room system and HMI were verified by a representative from the OEM/service provider (Guru Controls) on 02/04/2026 to identify the root cause. Upon verification, it was found that the batteries used for HMI time functionality were weak and that old version software was being used in the system’s HMI, which may have led to the HMI time going out of sync and subsequent communication failure between the HMI and the Kyoshi application. As per the service report, the OEM recommended replacing the batteries and upgrading the software to prevent any future non-conformances related to temperature data logging/capture. Although the non-conformance was identified, the Intermediate Walk-in Cold Room temperature was found to be within acceptable limits when verified at the last available instance, and no temperature excursion alarms occurred during the period when data was not logged in the system (Kyoshi application). Additionally, the daily MKT value on 06/02/2026 was 5.35 °C, which is within the acceptable range of 2 °C to 8 °C. Hence, there is no impact on the stored product/material and regulatory notification is not required. Since the non-conformance was a repeat occurrence, the previous CAPA was reassessed. It was found that the root cause had been identified as a loss of HMI time synchronization with the current date and time due to battery discharge or malfunction of the battery used for HMI time functionality. The CAPA for the previous non-conformance included replacement of the batteries for the respective cold room and incorporation of HMI battery verification instructions into the AMC/PM checklist for Blue Star make walk-in cold rooms. It was also identified that, although the CAPA recommends battery replacement, the batteries were not replaced as the batteries were found in good condition, hence, the action plan was not executed. |  |  |  |
| Analyze: Various investigation tools used to analyze the root cause like 6M, 5 Why, Personnel interview, Brainstorming etc. |  |  |  |
| 6 M Method (If Applicable): - Man: Not Applicable. - Machine: Not Applicable. - Measurement: Not Applicable. - Material: Not Applicable. - Method: Not Applicable. - Milieu (Environment): Not Applicable. Conclusion: Not Applicable |  |  |  |
| Note: No any 6 M question shall be deleted in the investigation, if any question is not applicable, then mention in ans. as “Not Applicable”. |  |  |  |
| 5 Why Approach (If Applicable): 1. Why was the temperature data not recorded/captured in the Intermediate Walk-in Cold Room system software (Kyoshi application)? Ans. Communication failure occurred between the Cold Room HMI and the Kyoshi application. 2. Why did the communication failure occur? Ans. The date and time on the Cold Room HMI were found to be out of synchronization with the current date and time. Due to this time mismatch, communication failure alarms were triggered, leading to disruption in data logging. 3. Why was the date/time out of synchronization with the current date and time on the HMI? Ans. The HMI uses coin-cell batteries (e.g., CR2032) to maintain the real-time clock (RTC) and SRAM (Static Random-Access Memory) during power interruptions. If the HMI remains powered off for extended periods, the internal battery becomes solely responsible for maintaining these functions, leading to accelerated battery discharge. Additionally, battery degradation, internal defects, or manufacturing quality issues may contribute to premature battery failure, resulting in loss of HMI time synchronization. 4. WHY: Not Applicable Ans. Not Applicable 5. WHY: Not Applicable Ans. Not Applicable Conclusion: Based on the 5-Why analysis, it is concluded that the date and time on the HMI were out of synchronization with the current date and time. This occurred because the HMI remained powered off for extended periods, during which the internal battery was solely responsible for maintaining the real-time clock, leading to accelerated battery discharge. Additionally, potential battery defects or quality issues may have contributed to premature battery failure and loss of time synchronization. Due to this time mismatch, temperature data was not captured in the Intermediate Walk-in Cold Room (E/PK/010). The issue was communicated to the software vendor, M/s Guru Control System, and based on their investigation, it was concluded that the existing battery was weak. Furthermore, it was identified that old version software was being used in the system (Refer Attachment IV). |  |  |  |
| Brainstorming: Not Applicable |  |  |  |
| Other Tool if Any: Not Applicable |  |  |  |
| Other Tool if Any: Not Applicable |  |  |  |
| Investigation Outcome: |  |  |  |
| On 05/02/2026, the HMI date and time stopped auto-synchronizing with the current date and time, resulting in communication failure between the HMI and the Kyoshi application in the Intermediate Walk-in Cold Room system. Due to this communication failure, the actual cold room temperature data displayed on the HMI was not recorded/captured in the system (Kyoshi application). Based on the preliminary investigation, the most probable root cause was identified as repeated communication failure alarms. A total of seven (07) communication failure alarms were recorded on 06/02/2026 between 02:08:50 hrs. and 14:10:02 hrs. during which temperature data was not captured in the system. It was further observed that the system failed to auto-update the date on the HMI. At approximately 15:30 hrs. on 06/02/2026, the HMI displayed the date and time as 05/02/2026, 18:30 hrs. (approx.). The HMI time was manually synchronized on 06/02/2026 at 15:37 hrs. against Work Order No. WO/PK/26-005, after which temperature data recording resumed in the Kyoshi application. Therefore, it is concluded that the time synchronization issue led to disruption of temperature data logging during the affected period. The non-conformance was communicated to the Original Equipment Manufacturer (OEM), M/s Guru Control System. As part of the detailed investigation, the cold room system and HMI were verified by the OEM/service provider on 02/04/2026. Upon verification, it was found that the batteries used for HMI time functionality were weak and that old version software was installed in the system’s HMI. These factors may have contributed to the HMI time going out of synchronization, leading to communication failure between the HMI and the Kyoshi application. As per the service report, the OEM recommended replacement of the batteries and upgrading of the software to prevent recurrence of such non-conformances. The investigation was conducted using the 5-Why methodology to identify the root cause. Based on the problem statement, event chronology, and 5-Why analysis, it was concluded that the HMI date and time were out of synchronization with the current date and time due to prolonged power-off conditions. During such periods, the internal battery is solely responsible for maintaining the real-time clock, leading to accelerated battery discharge. Additionally, battery defects or quality issues may have contributed to premature battery failure, resulting in loss of synchronization, communication failure, and subsequent failure of data recording in the system. The findings were further confirmed by the OEM, who confirmed that the existing battery was weak and that old version software was in use, which may have contributed to the failure in temperature data recording. Based on the nature of nonconformance, chronology of events, and a detailed investigation, further root cause categorization has been performed as per format no. SOP/DP/QA/008-F04 (Root Cause Categorization format). |  |  |  |
| Identified Root Cause/ Probable Cause: |  |  |  |
| Based on the detailed investigation, most probable root cause as mentioned below, Primary Root Cause Level 1: Equipment / Instrument. Secondary Root Cause Level 2: Electrical Failure. Third Root Cause Level 3: Not Applicable. |  |  |  |
| Impact Assessment (System/ Document/ Product/ Equipment/Patient safety/Past batches): |  |  |  |
| System: The observed non-conformance is related to the unavailability of temperature data in the system (Kyoshi application) for the intermediate walk-in cold room (E/PK/010) for a temporary period. During this period, the equipment operated normally, and the temperature remained within the range of 2 °C to 8 °C. After time synchronization in the HMI, temperature data logging resumed. Therefore, there is no impact on the system. Document: The observed non-conformance is related to the unavailability of temperature data in the system (Kyoshi application) for the intermediate walk-in cold room (E/PK/010) for a temporary period. Relevant documents, including the Material Storage and Removal Log of the Intermediate Walk-in Cold Room (SOP/DP/PK/013/F03-R01), Alarm Report, and Audit Trail Report, were verified, and no abnormalities were observed. Temperature data from 06/02/2026, 03:47 to 06/02/2026, 15:31, were not available in the system data log, however, the cold room’s daily MKT value report for 06/02/2026 was available. Hence, there is no impact on the document. Product: The observed non-conformance is related to the unavailability of temperature data in the system (Kyoshi application) for the intermediate walk-in cold room (E/PK/010) for a temporary period. Semi-finished and finished product batches, as listed in the ‘List of Product/Material Stored in the Intermediate Walk-in Cold Room’ document, are stored in the cold room during this period. No temperature excursions (low or high temperature alarms) alarms occurred, and the daily MKT value on 06/02/2026 was 5.35 °C, which is within the acceptable range of 2 °C to 8 °C. Additionally, 6-month accelerated stability data at 25°C and 60% relative humidity are available for products stored at 2°C to 8°C. Therefore, there was no impact on the stored products or materials. Equipment: The non-conformance occurred due to the HMI date and time being out of synchronization with the actual date and time because of a weak battery. During this period, only data recording was stopped, the equipment continued to operate, and temperature was maintained. Hence, there was no impact on the equipment. Patient Safety / Past Batches: At the time of the non-conformance, product batches/materials were stored in the intermediate walk-in cold room (E/PK/010). No temperature excursion alarms were observed, and the daily MKT value remained within limits. This confirms that the cold room temperature was within acceptable limits. Therefore, there was no impact on patient safety or past batches. |  |  |  |
| Improve: Improve section covers the corrective actions Following checkpoint shall be considered while finalizing the corrective actions, - Were specific corrective Actions identified (including applicable Immediate Actions) to remediate the current issue or situation such that the associated system was returned to a state of control/compliance? - Were specific Corrective Actions identified for each root cause /substantiated probable root cause, as applicable? - Was the Corrective action assigned a unique number, responsible person and due date so it can be tracked? Does the action describe what will be the expected outcome they can be verified? - Was Effectiveness Verification required or not and the rationale for either documented based within the quality impact of the deviation after the investigation concluded? - Are the identified corrective actions achievable based on the information provided? |  |  |  |
| Corrective Action: |  |  |  |
| The non-conformance is related to temperature data not being captured in the system for the Intermediate Walk-in Cold Room (Equipment ID: E/PK/010) for a temporary period. The investigation was conducted using the 5-Why methodology. Based on the detailed investigation, it was concluded that the non-conformance occurred because the HMI remained powered off for extended periods, during which the internal battery was solely responsible for maintaining the real-time clock. This led to accelerated battery discharge and subsequent loss of time synchronization on the HMI, resulting in failure to capture temperature data in the system. Based on the identified non-conformance, the following corrective actions were taken: - Work Order No. WO/PK/26-005 was initiated for rectification of the HMI time asynchronization issue. - The date and time were corrected through the HMI internal settings for Equipment ID: E/PK/010. - The incident was escalated to the service provider, M/s Guru Control System, for further investigation and support. - Awareness training on deviation handling was imparted to the concerned personnel. A CAPA form shall be initiated to track the corrective actions with a unique identification number. Responsibilities and target completion dates shall be defined in the CAPA form. The effectiveness and implementation of the corrective actions shall be verified and documented in the CAPA form. |  |  |  |
| Control: Control section covers the preventive actions Following checkpoint shall be considered while finalizing the preventive actions, - Were specific Preventive Actions identified for each root cause /substantiated probable root cause as applicable? - Was the Preventive Action linked the classification of the root cause and explanation given for how it will prevent recurrence? - Was the Preventive action assigned a unique number, responsible person and due date so it can be tracked? Does the action describe and expected outcome can be verified? - Was Effectiveness Verification required or not and the rationale for either documented based on the quality impact of the incident after the investigation concluded? - Was an Interim Plan needed to ensure a state the control while the Preventive Actions were implemented? - Was rationale provided when no Preventive Action were identified? - Does the Final Comments section include rotational to support the conclusion of the investigation and CAPA. - Was each of the impact assessment fields completed correctly - Regulatory Impact, Regulatory notification Product Quality, Validation, Stability, Market/Clinical? - Does the Recommended Lot disposition match the conclusions of the investigation and Impact assessment? - Does the Conclusion include final decision and rationale other regulatory notification required? - CAPA required to release material or batches have been verified to be complete and closed prior to material or batch disposition. Any related CAPA that remain open, but have no impact on the material or batch release and a documented explanation is provided. - Does the Conclusion include a brief summary of root cause, final scope/impact including rationale for product/material disposition, impact assessment and relevant lot/material/equipment details? - Are the identified preventive actions achievable based on the information provided? |  |  |  |
| Preventive Action: |  |  |  |
| The investigation was carried out using the 5-Why methodology. During the investigation, the root cause was identified: the date and time on the HMI were out of synchronization with the actual date and time. This occurred because the HMI remained powered off for extended periods, leaving the internal battery solely responsible for maintaining these functions, which led to accelerated battery discharge. Based on the investigation outcome and the OEM’s service report, the following preventive actions have been identified: - HMI batteries shall be replaced for all Blue Star walk-in cold rooms (E/PK/010, E/WH/020, E/WH/022, E/WH/027 & E/QC/019) and walk-in deep freezers (E/WH/021 & E/WH/023). - Upgraded software shall be installed for all Blue Star walk-in cold rooms (E/PK/010, E/WH/020, E/WH/022, E/WH/027 & E/QC/019) and walk-in deep freezers (E/WH/021 & E/WH/023). A CAPA form will be initiated to track the preventive actions with a unique identification number. Responsibilities and target completion dates for the identified preventive actions will be assigned in the CAPA form. The implementation outcome of the preventive actions will be verified and documented within the CAPA form. The non-conformance is categorized as major, and the identified actions relate specifically to the replacement of HMI batteries and software upgradation in Blue Star walk-in cold rooms and walk-in freezer, therefore, effectiveness monitoring is not required. The preventive actions will be implemented following CAPA approval. Currently, the intermediate walk-in cold is in healthy operating condition, and daily data monitoring procedure is in place, therefore, interim controls are not required. |  |  |  |
| Document Reviewed: |  |  |  |
| - SOP for operation and cleaning procedure for intermediate cold room (SOP/DP/PK/013). |  |  |  |
| List of attachment (If applicable): |  |  |  |
| - Attachment No. I: Photocopy of Work Order Form - Attachment No. II: Audit Trail Report - Attachment No. III: Preliminary Investigation - Attachment No. IV: Service Provider Justification Report - Attachment No. V: Root Cause Categorization |  |  |  |

| Prepared By (Sign/Date) | Reviewed By (Sign/Date) | Reviewed By (Sign/Date) | Reviewed By (Sign/Date) | Reviewed By (Sign/Date) | Reviewed By (Sign/Date) | Reviewed By (Sign/Date) | Approved By QA (Sign/Date) |
| --- | --- | --- | --- | --- | --- | --- | --- |