| Date: | 08/04/2026 | Deviation No. | DEV/PK/25/002 |
| --- | --- | --- | --- |
| Investigation tool used: 6M 5 Why Brainstorming |  |  |  |
| Other Tools (If any): |  |  |  |
| Not applicable |  |  |  |
| Define: |  |  |  |
| Following checks shall be considered while writing the “Define” section. - Clearly define what happens actually. - Explain what is different than expected. - Mention the location where the deviation has occurred. - Date/time of deviation occurrence and date/time of detection. - Mention the name of personnel who is involved in the deviation. - Mention initial scope of deviation (i.e., impacted product/Material/Equipment/System/Batches/etc.) On dated 04/03/2026 at approximately 16:20 hrs, while reviewing the temperature data logs/reports of Intermediate Walk-in Cold Room (Temperature range: 2°C to 8°C), equipment ID: E/PK/010 at packing office (GF-109), it was observed by a packing colleague (employee ID: 598) that the temperature data was not recorded in the system (Kyoshi application) for the period of 06/02/2026, 03:47 to 06/02/2026, 15:31. As per the site SOP No. SOP/DP/PK/013 (Operation and Cleaning of Intermediate Walk-in Cold Room) point no. 7.4.2 stated that “The intermediate cold room temperature will be recorded in the system through Kyoshi software”, however the temperature data was not recorded for the period of 06/02/2026, 03:47 to 06/02/2026, 15:31, which is against the SOP. The nonconformance occurred on the ground floor (GF-122) at packing area. At the time of nonconformance the product batches were stored in intermediate walk-in cold. |  |  |  |
| Details Investigation: |  |  |  |
| Measure: Following checks shall be considered while writing the “Measure” section. - Does the summary provide relevant facts and data/information that was reviewed including: environment, process/product history, personnel info (title and job title), controls limits, etc. - Is a summary of the analysis of the factors and data provided? - Is a conclusion statement of the analysis and review provided? - If there were Regulatory Notification, were details provided? - Is the report written in a logical flow and easily understood by the reader? The intermediate walk-in cold room (Equipment ID: E/PK/010) is installed in the Packing area (GF-122) for storage of semi-finished goods (temperature range: 2°C to 8°C). SOP No. SOP//DP/PK/013 (Operation and Cleaning of intermediate Walk-in Cold Room) is in place for cleaning and operation of intermediate Walk-in Cold Room. As per this SOP the intermediate cold room temperature data shall be recorded/captured in the system through Kyoshi software. However, on dated 06/02/2026 the temperature data was not recorded/captured in the system from 03:47 to 15:31. On dated 06/02/2026 at approximately 10:30 hrs , it was identified by packing colleague (Employee ID: xxx) that the Intermediate Walk-in Cold Room HMI time was found mismatched with the current time, hence a work order request (Breakdown no. WO/PK/25-007) was initiated to inline the HMI time with current time. The HMI time was updated with the current time by IT colleague (Employee ID: xxx) at approximately 15:40 hrs against the work order. There after the temperature data recording was started in the system (Kyoshi application). The work order was then received by the packing colleague (Employee ID:598) on 06/02/2026 to complete the evaluation and verification of user part. A part of the verification the HMI time and temperature data recording were kept under monitoring till 04/03/2026. After changing the date/time in HMI, the temperature data was thoroughly monitored, and no data was found missing till date. and no time mismatch was observed in the HMI. And upon reviewing the temperature and alarm reports on 04/03/2026, the temperature data was found recorded/captured from 06/03/2026 15:32 hrs to 04/03/2026 09:23, however, the temperature data during the time asynchronous period (06/02/2026, 03:47 to 06/02/2026, 15:31) was not recorded in the reports, and it is also found that the communication failure alarm was generated during this period. Due to this time asynchronous, the system may have failed to capture the temperature data from 06/02/2026, 03:47 to 06/02/2026, 15:31. During this period semi-finished product batches were stored in the intermediate Walkin cold room. A deviation was initiated on 05/03/2026 at 13:01 to identify the root cause for identified non-conformance. Although non-conformance was identified, the intermediate walk-in cold room temperature were found within acceptable limits when they were verified at the last instant and no temperature excursion alarms occurred during this period. Hence there is no impact on stored materials and regulatory notification is not required. |  |  |  |
| Analyze: Various investigation tools used to analyze the root cause like 6M, 5 Why, Personnel interview, Brainstorming etc. |  |  |  |
| 6 M Method (If Applicable): - Man: Not Applicable. - Machine: Not Applicable. - Measurement: Not Applicable. - Material: Not Applicable. - Method: Not Applicable. - Milieu (Environment): Not Applicable. Conclusion: Not Applicable |  |  |  |
| Note: No 6 M question shall be deleted in the investigation, if any question is not applicable, then mention in ans. as “Not Applicable”. |  |  |  |
| 5 Why Approach (If Applicable): 1. Why was the temperature data not recorded/captured in intermediate walk-in cold room systems software (Kyoshi application? Ans. Communication failure between HMI and system. 2. Why was the communication failure happened? Ans. The date/time was found out of synchronize with current time date and time on cold room HMI. Due to this time asynchronous, communication failure alarm occurred. 3. Why the date/time was out of synchronize with current time date and time on HMI? Ans. In the HMIs, coin-cell batteries (such as CR2032) are used to maintain the real-time clock (RTC) and SRAM (Static Random-Access Memory) data during power outages. If the HMI remains powered off for extended periods, the internal battery is solely responsible for maintaining these functions, which leads to accelerated battery discharge. Additionally, some batteries may suffer from internal defects or manufacturing quality issues, further contributing to early battery failure and loss of HMI time synchronization. 4. WHY: Not Applicable Ans. Not Applicable 5. WHY: Not Applicable Ans. Not Applicable Conclusion: Based on the 5 why tool, it is concluded that the date and time was out of synchronize with current time date and time on HMI. The HMI remains powered off for extended periods, the internal battery is solely responsible for maintaining these functions, which leads to accelerated battery discharge. Additionally, some batteries may suffer from internal defects or manufacturing quality issues, further contributing to early battery failure and loss of synchronization. Due to time synchronization leading to data not captured in the Intermediate Walk-in Cold Room (E/PK/010). The issue was communicated with software vendor “M/s Guru Control System” and based on the vendor investigation it was declared that the existing battery was found weak and additionally it was also found that an old software was being used in the system (Refer attachment I). |  |  |  |
| Brainstorming: Not Applicable |  |  |  |
| Other Tool if Any: Not Applicable |  |  |  |
| Other Tool if Any: Not Applicable |  |  |  |
| Investigation Outcome: |  |  |  |
| The nonconformance is related to the temperature data that was not recorded/captured in the intermediate walk-in cold room (Equipment ID: E/PK/010) systems software. The investigation was driven through the 5 Why to identify the root cause. Based on the problem statement and chronology of event and 5 Why analysis, the date and time was out of synchronize with current time date and time on HMI due to the HMI remains powered off for extended periods, the internal battery is solely responsible for maintaining these functions, which leads to accelerated battery discharge. Additionally, some batteries may suffer from internal defects or manufacturing quality issues, further contributing to early battery failure and loss of synchronization leading to communication failure and followed by data not recorded in the system. The identified nonconformance was communicated with software vendor “M/s Guru Control System” and based on the vendor investigation it was confirmed that the existing battery was found weak and additionally it was also found that an old software was being used in the system (Refer attachment I) which may be leads to data not recorded. Based on the nature of nonconformance, a chronology of events, and a details investigation, further root cause categorization has been done as per format for “Root Cause Categorization” SOP/DP/QA/008-F04 |  |  |  |
| Identified Root Cause/ Probable Cause: |  |  |  |
| Based on the detailed investigation, most probable root cause as mentioned below, Primary Root Cause Level 1: Equipment / Instrument. Secondary Root Cause Level 2: Not Applicable. Third Root Cause Level 3: Not Applicable. |  |  |  |
| Impact Assessment (System/ Document/ Product/ Equipment/Patient safety/Past batches): |  |  |  |
| System: The nonconformance occurred due to the out of date time synchronization with the actual HMI due to the battery discharge. Data recording was stopped, but equipment operation was running smoothly and temperature was maintained. There is no impact on the system. Document: The nonconformance about the temperature data was not recorded for intermediate walk-in cold room (equipment ID: E/PK/010) for a specific period. The equipment routine used documents have been verified, and no abnormality was observed. Hence, there is no impact on the document. Product: The nonconformance is related to the data was not recorded for intermediate walk-in cold room (equipment ID: E/PK/010) for a specific period. At the time of nonconformance, insulin product batches were stored in intermediate walk-in cold room (equipment ID: E/PK/010. While the non-conformance was observed, the intermediate walk-in cold room (equipment ID: E/PK/010) daily MKT value was found within the acceptance limit. There were not any alarms generated for temperature excursion during this specific time period. Hence there is no impact on the product. Equipment: The nonconformance occurred due to the out of date time synchronization with the actual HMI due to the battery discharge. During the nonconformance, only data recording was stopped, but equipment operation remained operational and temperature was maintained. There is no impact on the system. Patient safety / Past Batches: At the time of nonconformance, the product batches/materials were stored in the intermediate walk-in cold room (equipment ID: E/PK/010). While the non-conformance was observed, the intermediate walk-in cold room (equipment ID: E/PK/010) temperature was within the acceptance limit. Hence there is no impact on the patient safety/past batches. |  |  |  |
| Improve: Improve section covers the corrective actions Following checkpoint shall be considered while finalizing the corrective actions, - Were specific corrective Actions identified (including applicable Immediate Actions) to remediate the current issue or situation such that the associated system was returned to a state of control/compliance? - Were specific Corrective Actions identified for each root cause /substantiated probable root cause, as applicable? - Was the Corrective action assigned a unique number, responsible person and due date so it can be tracked? Does the action describe what will be the expected outcome they can be verified? - Was Effectiveness Verification required or not and the rationale for either documented based within the quality impact of the deviation after the investigation concluded? - Are the identified corrective actions achievable based on the information provided? |  |  |  |
| Corrective Action: |  |  |  |
| The nonconformance is related to the temperature data that was not captured in the system for intermediate walk-in cold room (Equipment ID: E/PK/010) for a specific time. After identification of the nonconformance below actions were taken, - The breakdown no. BD/PK/25-001 was initiated for rectification of HMI time asynchronous issue. - The date and time have been changed through HMI internal setting for E/PK/010. - The incident has been escalated to the service provider, M/s Guru Control System. - Awareness training on deviation has been imparted to concerned personnel. The investigation has been driven by 5 Why methodology. Based on the detailed investigation, it has been concluded that the nonconformance occurred because the HMI remained powered off for extended periods, the internal battery was solely responsible for maintaining these functions, which led to accelerated battery discharge. Due to the battery discharge, the time synchronization activity was not properly done on the HMI. Hence, it leads to data not captured in the intermediate walk-in cold room (Equipment ID: E/PK/010). Following corrective action has been identified based on the investigation outcome, - The intermediate walk-in cold room (E/PK/010) HMI batteries to be replaced. CAPA form shall be initiated to track the corrective action with a unique number. Responsibilities and target completion date shall be mentioned in the CAPA form. The implementation outcome of corrective action shall be verified in CAPA form. |  |  |  |
| Control: Control section covers the preventive actions Following checkpoint shall be considered while finalizing the preventive actions, - Were specific Preventive Actions identified for each root cause /substantiated probable root cause as applicable? - Was the Preventive Action linked the classification of the root cause and explanation given for how it will prevent recurrence? - Was the Preventive action assigned a unique number, responsible person and due date so it can be tracked? Does the action describe and expected outcome can be verified? - Was Effectiveness Verification required or not and the rationale for either documented based on the quality impact of the incident after the investigation concluded? - Was an Interim Plan needed to ensure a state the control while the Preventive Actions were implemented? - Was rationale provided when no Preventive Action were identified? - Does the Final Comments section include rotational to support the conclusion of the investigation and CAPA. - Was each of the impact assessment fields completed correctly - Regulatory Impact, Regulatory notification Product Quality, Validation, Stability, Market/Clinical? - Does the Recommended Lot disposition match the conclusions of the investigation and Impact assessment? - Does the Conclusion include final decision and rationale other regulatory notification required? - CAPA required to release material or batches have been verified to be complete and closed prior to material or batch disposition. Any related CAPA that remain open, but have no impact on the material or batch release and a documented explanation is provided. - Does the Conclusion include a brief summary of root cause, final scope/impact including rationale for product/material disposition, impact assessment and relevant lot/material/equipment details? - Are the identified preventive actions achievable based on the information provided? |  |  |  |
| Preventive Action: |  |  |  |
| The investigation has been carried out through 5 Why methodology. During the investigation root cause has been identified as the date and time was out of synchronize with current time date and time on HMI due to the HMI remains powered off for extended periods, the internal battery is solely responsible for maintaining these functions, which leads to accelerated battery discharge. As per the outcome of the investigation report and service provider justification report following preventive action has been identified. - Other blue star walk-in cold rooms (E/WH/022, E/WH/027 & E/QC/019) and walk-in deep freezers (E/WH/023) batteries to be checked and if required replacement to be done. - Software to be updated for all blue star walk-in cold rooms and deep freezers. CAPA form shall be initiated to track the preventive action with unique numbering. Responsibility and target completion date of identified preventive actions shall be assigned in CAPA form. The implementation Outcome of preventive action shall be verified in CAPA form. The nonconformance is isolated and categorized as major and identified actions are related to the replace of the HMI batteries of blue star make walk-in cold room and walk-in freezer room HMIs, hence effectiveness monitoring shall not require. The preventive action shall be implemented after approval of CAPA, currently both equipment’s operation is healthy condition and data monitoring on daily basis procedure is in place, hence interim control not required. The occurred deviation is related to data unavailability of intermediate walk-in cold room and root cause identified as equipment /instrument failure, hence no impact on product quality, validation, stability, market/clinical and testing/release of product and no regulatory impact. The nonconformance is related to the data unavailability of intermediate walk-in cold room, there is not any impact to the products/materials, hence material or batch disposition impact not anticipated. |  |  |  |
| Document Reviewed: |  |  |  |
| - SOP for operation and cleaning procedure for intermediate cold room. |  |  |  |
| List of attachment (If applicable): |  |  |  |
| - Attachment No. I: Photocopy of Breakdown Form. - Attachment No. II: Preliminary Investigation - Attachment No. IV: Training Documentation Form - Attachment No. V: Service Provider Justification Report - Attachment No. VI: Root Cause Categorization |  |  |  |

| Prepared By (Sign/Date) | Reviewed By (Sign/Date) | Reviewed By (Sign/Date) | Reviewed By (Sign/Date) | Reviewed By (Sign/Date) | Reviewed By (Sign/Date) | Reviewed By (Sign/Date) | Approved By QA (Sign/Date) |
| --- | --- | --- | --- | --- | --- | --- | --- |