## Summary
Deviation Portfolio — At-a-Glance
Dev ID | Dept | Category | Non-Conformance Type | Root Cause L1 | Repeat? | # Corrective Actions
DEV/EU/25/004 | EU | Minor | Utility interruption during SIP requalification | Equipment / Instrument | No | 3
DEV/HR/23/001 | HR | Minor | Documentation / procedural non-compliance | Human Performance | ⚑ YES | 3
DEV/HR/23/002 | HR | Minor | Training/Documentation nonconformance | Human Performance | No | 2
DEV/IT/25/002 | QC | Minor | Data integrity / audit trail data loss due to backup overwrite | Human error / Method / Procedure | No | 1
DEV/PK/25/002 | PK | Major | Data logging / temperature monitoring system failure due to HMI time synchronization loss and communication failure | Equipment / Instrument | ⚑ YES | 5
DEV/PR/24/016 | PR | Minor | Document control / uncontrolled missing controlled copy of approved vendor list (AVL) and related SOP in production department | Human Error | ⚑ YES | 3
DEV/PR/25/003 | PR | Minor | Audit trail data not visible / audit trail generation failure in SCADA system | Equipment | ⚑ YES | 2
DEV/PR/25/004 | PR | Minor | Equipment/Instrument malfunction leading to sterilizer cycle temperature overshoot and cycle abort | Equipment/Instrument | ⚑ YES | 2
DEV/PR/25/008 | PR | Minor | Incorrect cycle selection / procedural non-adherence (human error) | Human Performance | No | 3
DEV/QC/25/002 | QC | Minor | Analytical instrument calibration OOC during HPLC wavelength accuracy check due to pressure drop/air bubble in pump | Instrument error | No | 2
DEV-WH-24-004 | ? | Minor | ? | Equipment / Instrument | No | 3
DEV/WH/25/003 | WH | Minor | Document control / Good Documentation Practices breach | Human error | No | 4

## RCA → Actions
RCA to Corrective / Preventive Action Mapping
Dev ID | Department | Category | Non-Conformance Type | Root Cause (L1) | Root Cause (L2) | Root Cause (L3) | Action Type | Action Category | Action Description | ⚑ Flag
DEV/EU/25/004 | EU | Minor | Utility interruption during SIP requalification | Equipment / Instrument | Mechanical | Gasket deterioration / seal failure | CORRECTIVE: immediate | awareness_training | Provide awareness training to concerned personnel that pure steam generation must not be manually stopped during an active SIP cycle unless absolutely necessary and escalated appropriately. | ⚠ Training Only
|  |  |  | Equipment / Instrument | Mechanical | Gasket deterioration / seal failure | CORRECTIVE: corrective | equipment_repair | Replace the damaged silicone TC gasket with a PTFE gasket to improve thermal and mechanical durability under steam sterilization conditions. | ✔ Action
|  |  |  | Equipment / Instrument | Mechanical | Gasket deterioration / seal failure | CORRECTIVE: corrective | equipment_repair | Replace the WFI system gasket with a new PTFE gasket. | ✔ Action
|  |  |  | Equipment / Instrument | Mechanical | Gasket deterioration / seal failure | PREVENTIVE: ? | pm_checklist_revision | Verify all gaskets in the pure steam generation and distribution system during scheduled PM activities and replace any gasket showing wear, deformation, or degradation based on physical condition rather than service life. | ✔ Action
|  |  |  | Equipment / Instrument | Mechanical | Gasket deterioration / seal failure | PREVENTIVE: ? | pm_checklist_revision | Revise PM checklists PMC/EU/006 and PMC/EU/007 to include specific gasket condition inspection points and replacement criteria. | ✔ Action
|  |  |  | Equipment / Instrument | Mechanical | Gasket deterioration / seal failure | PREVENTIVE: ? | other | Initiate a CAPA form with unique number, responsible person, and target completion date to track the preventive actions. | ✔ Action
DEV/HR/23/001 | HR | Minor | Documentation / procedural non-compliance | Human Performance | Omission Error | Responsibility boundaries ineffective | CORRECTIVE: immediate | document_reissuance | Revise the Corporate Communication PPT (PPT/020-R00) through change control to align with approved SOP format. | ✔ Action
|  |  |  | Human Performance | Omission Error | Responsibility boundaries ineffective | CORRECTIVE: corrective | awareness_training | Impart awareness training to all concerned personnel regarding Deviation DEV/HR/23/001 and PPT SOP requirements. | ⚠ Training Only
|  |  |  | Human Performance | Omission Error | Responsibility boundaries ineffective | CORRECTIVE: corrective | document_reissuance | Revise the listed PPTs (PPT/001-R00, PPT/006-R00, PPT/011-R00, PPT/020-R00) to align with SOP/DP/QA/033/F01. | ✔ Action
|  |  |  | Human Performance | Omission Error | Responsibility boundaries ineffective | PREVENTIVE: ? | awareness_training | Impart awareness training to concerned personnel on SOP/DP/QA/033-R00 for handling PPTs. | ✔ Action
|  |  |  | Human Performance | Omission Error | Responsibility boundaries ineffective | PREVENTIVE: ? | sop_revision | Revise SOP/DP/QA/033-R00 to require QA personnel review before approval of any PPT. | ✔ Action
DEV/HR/23/002 | HR | Minor | Training/Documentation nonconformance | Human Performance | Omission Error | Know but confused – Responsibility boundaries non-effective | CORRECTIVE: immediate | document_reissuance | Issue an additional page 04 of 05 of the New Employee Orientation form and complete drug product SCM-Warehouse/PPIC orientation for the affected employee(s). | ✔ Action
|  |  |  | Human Performance | Omission Error | Know but confused – Responsibility boundaries non-effective | CORRECTIVE: corrective | awareness_training | Conduct awareness training on the deviation for concerned persons. | ⚠ Training Only
|  |  |  | Human Performance | Omission Error | Know but confused – Responsibility boundaries non-effective | PREVENTIVE: ? | awareness_training | Prepare a PPT on the usage of files for new employees and conduct training/awareness to prevent misidentification. | ✔ Action
DEV/IT/25/002 | QC | Minor | Data integrity / audit trail data loss due to backup overwrite | Human error | Omission error | Responsibility boundaries ineffective | CORRECTIVE: immediate | other | Identify the equipment/instrument applicable for data archival and prepare the list. | ✔ Action
|  |  |  | Method / Procedure | Document lacking details | Vendor instruction not embedded in SOP/protocol | CORRECTIVE: immediate | other | Identify the equipment/instrument applicable for data archival and prepare the list. | ✔ Action
|  |  |  | Human error | Omission error | Responsibility boundaries ineffective | PREVENTIVE: ? | sop_revision | Revise SOP/DP/IT/001 for procedural clarity so that every backup file generated before archival is renamed with the date to prevent overwrite of current backup data. | ✔ Action
|  |  |  | Method / Procedure | Document lacking details | Vendor instruction not embedded in SOP/protocol | PREVENTIVE: ? | sop_revision | Revise SOP/DP/IT/001 for procedural clarity so that every backup file generated before archival is renamed with the date to prevent overwrite of current backup data. | ✔ Action
DEV/PK/25/002 | PK | Major | Data logging / temperature monitoring system failure due to HMI time synchronization loss and communication failure | Equipment / Instrument | Electrical Failure | Not Applicable | CORRECTIVE: immediate | other | Work Order No. WO/PK/26-005 was initiated to rectify the HMI time asynchronization issue. | ✔ Action
|  |  |  | Equipment / Instrument | Electrical Failure | Not Applicable | CORRECTIVE: immediate | equipment_repair | The date and time were corrected through the HMI internal settings for Equipment ID E/PK/010. | ✔ Action
|  |  |  | Equipment / Instrument | Electrical Failure | Not Applicable | CORRECTIVE: immediate | vendor_escalation | The incident was escalated to the service provider, M/s Guru Control System, for further investigation and support. | ✔ Action
|  |  |  | Equipment / Instrument | Electrical Failure | Not Applicable | CORRECTIVE: immediate | awareness_training | Awareness training on deviation handling was imparted to the concerned personnel. | ⚠ Training Only
|  |  |  | Equipment / Instrument | Electrical Failure | Not Applicable | CORRECTIVE: corrective | other | Initiate and track a CAPA form with unique identification, responsibilities, and target completion dates for the corrective actions. | ✔ Action
|  |  |  | Equipment / Instrument | Electrical Failure | Not Applicable | PREVENTIVE: ? | equipment_replacement | Replace HMI batteries for all Blue Star walk-in cold rooms (E/PK/010, E/WH/020, E/WH/022, E/WH/027, E/QC/019) and walk-in deep freezers (E/WH/021, E/WH/023). | ✔ Action
|  |  |  | Equipment / Instrument | Electrical Failure | Not Applicable | PREVENTIVE: ? | software_update | Install upgraded software for all Blue Star walk-in cold rooms (E/PK/010, E/WH/020, E/WH/022, E/WH/027, E/QC/019) and walk-in deep freezers (E/WH/021, E/WH/023). | ✔ Action
DEV/PR/24/016 | PR | Minor | Document control / uncontrolled missing controlled copy of approved vendor list (AVL) and related SOP in production department | Human Error | Omission Error | Responsibility boundaries ineffective | CORRECTIVE: immediate | other | Inform Production Head and QA Head and initiate deviation DEV/PR/24/016. | ✔ Action
|  |  |  | Human Error | Omission Error | Responsibility boundaries ineffective | CORRECTIVE: corrective | awareness_training | Provide awareness training to all concerned personnel regarding deviation DEV/PR/24/016 and document handling expectations. | ⚠ Training Only
|  |  |  | Human Error | Omission Error | Responsibility boundaries ineffective | CORRECTIVE: corrective | document_reissuance | Issue an additional controlled copy of SOP for cleaning validation (SOP/DP/QA/046-R00 as written in report) from QA. | ✔ Action
|  |  |  | Human Error | Omission Error | Responsibility boundaries ineffective | PREVENTIVE: ? | infrastructure_modification | Keep document files in a locker with lock-and-key provision to prevent uncontrolled access and misplacement. | ✔ Action
DEV/PR/25/003 | PR | Minor | Audit trail data not visible / audit trail generation failure in SCADA system | Equipment | Computer system | Scheduled cleanup task failure / database table full (IFIXALARMS) | CORRECTIVE: immediate | other | Information given to production head and QA head after identification of the nonconformance and deviation DEV/PR/25/001 initiated. | ✔ Action
|  |  |  | Equipment | Computer system | Scheduled cleanup task failure / database table full (IFIXALARMS) | CORRECTIVE: immediate | vendor_escalation | Work order WO/PR/25-024 raised to IT/OEM and issue resolved to restore audit trail functionality. | ✔ Action
|  |  |  | Equipment | Computer system | Scheduled cleanup task failure / database table full (IFIXALARMS) | PREVENTIVE: ? | no_action_required | No standalone CAPA proposed because OEM resolved the issue, recurrence is not anticipated, and audit trail review/backup-restoration procedures are already in place. | ℹ No CAPA
DEV/PR/25/004 | PR | Minor | Equipment/Instrument malfunction leading to sterilizer cycle temperature overshoot and cycle abort | Equipment/Instrument | Mechanical failure | Not Applicable | CORRECTIVE: immediate | equipment_repair | Inform production head and QA head; raise breakdown BD/PR/25-018 to rectify the issue; perform repairs and verification trials. | ✔ Action
|  |  |  | Equipment/Instrument | Mechanical failure | Not Applicable | CORRECTIVE: corrective | pm_checklist_revision | Revise PM checklist to include operation verification of jacket steam inlet valve (101). | ✔ Action
|  |  |  | Equipment/Instrument | Mechanical failure | Not Applicable | PREVENTIVE: ? | pm_checklist_revision | Revise PM checklist to include operation verification of jacket steam inlet valve (101) during PM. | ✔ Action
DEV/PR/25/008 | PR | Minor | Incorrect cycle selection / procedural non-adherence (human error) | Human Performance | Omission error | Responsibility boundaries ineffective | CORRECTIVE: immediate | equipment_repair | Re-perform load cell calibration against Work Order WO/PR/26-012. | ✔ Action
|  |  |  | Human Performance | Omission error | Responsibility boundaries ineffective | CORRECTIVE: immediate | document_reissuance | Generate and verify recalibration report with correct data capture. | ✔ Action
|  |  |  | Human Performance | Omission error | Responsibility boundaries ineffective | CORRECTIVE: immediate | awareness_training | Provide awareness training to concerned operators on correct cycle selection and SOP compliance. | ⚠ Training Only
|  |  |  | Human Performance | Omission error | Responsibility boundaries ineffective | PREVENTIVE: ? | no_action_required | No standalone CAPA proposed because the event was a first-time isolated human error with no equipment, system, or procedure deficiency identified. | ℹ No CAPA
DEV/QC/25/002 | QC | Minor | Analytical instrument calibration OOC during HPLC wavelength accuracy check due to pressure drop/air bubble in pump | Instrument error | Mechanical Failure | Air bubble trapped in pump/check-inlet tubing causing pressure drop | CORRECTIVE: immediate | equipment_repair | Purged the HPLC system I/QC/021, verified stabilized pressure, repeated the wavelength accuracy test, and used the repeat sequence 021_0102 for reporting. | ✔ Action
|  |  |  | Instrument error | Mechanical Failure | Air bubble trapped in pump/check-inlet tubing causing pressure drop | CORRECTIVE: immediate | document_reissuance | Invalidated chromatographic data from sequences 021_0100 and 021_0101 due to the affected pressure-drop event. | ✔ Action
|  |  |  | Instrument error | Mechanical Failure | Air bubble trapped in pump/check-inlet tubing causing pressure drop | PREVENTIVE: ? | awareness_training | Provide awareness training to all concerned personnel to purge the HPLC system properly, verify system pressure before analysis, and start sequences only after stabilized pressure is achieved. | ✔ Action
DEV-WH-24-004 | ? | Minor | ? | Equipment / Instrument | Electrical Failure | Damaged cable insulation causing short circuit and MCB tripping | CORRECTIVE: immediate | other | Inform Warehouse Head and Quality Assurance Head of the nonconformance. | ✔ Action
|  |  |  | Equipment / Instrument | Electrical Failure | Damaged cable insulation causing short circuit and MCB tripping | CORRECTIVE: corrective | awareness_training | Impart awareness training on deviation DEV/WH/24/004 to concerned persons. | ⚠ Training Only
|  |  |  | Equipment / Instrument | Electrical Failure | Damaged cable insulation causing short circuit and MCB tripping | CORRECTIVE: corrective | awareness_training | Impart awareness training on Breakdown and Work Order Procedure SOP/DP/EU/003 to concerned persons. | ⚠ Training Only
|  |  |  | Equipment / Instrument | Electrical Failure | Damaged cable insulation causing short circuit and MCB tripping | PREVENTIVE: ? | equipment_repair | Check the cable coming from the UPS supply for proper condition. | ✔ Action
|  |  |  | Equipment / Instrument | Electrical Failure | Damaged cable insulation causing short circuit and MCB tripping | PREVENTIVE: ? | infrastructure_modification | Provide separate UPS supply for each cold room. | ✔ Action
DEV/WH/25/003 | WH | Minor | Document control / Good Documentation Practices breach | Human error | Omission error | Responsibility boundaries ineffective | CORRECTIVE: immediate | document_reissuance | Sent impacted documents (Retest and Expiry Material Checklist June 2025, SAP report templates, weighing slips, JD copy) to DS facility and destroyed the checklist as per approved procedure. | ✔ Action
|  |  |  | Human error | Omission error | Responsibility boundaries ineffective | CORRECTIVE: immediate | awareness_training | Imparted awareness training to the Warehouse HOD on Good Documentation Practices and Document and Data Control. | ⚠ Training Only
|  |  |  | Human error | Omission error | Responsibility boundaries ineffective | CORRECTIVE: immediate | awareness_training | Provided awareness training on Good Documentation Practices to concerned persons. | ⚠ Training Only
|  |  |  | Human error | Omission error | Responsibility boundaries ineffective | CORRECTIVE: corrective | awareness_training | Prepare and deliver a presentation/training on the procedure for handling documents between DS and DP units. | ✔ Action
|  |  |  | Human error | Omission error | Responsibility boundaries ineffective | PREVENTIVE: ? | no_action_required | No separate preventive action proposed; issue was attributed to a one-time human omission and the immediate awareness training was considered sufficient. | ℹ No CAPA

## Repeat Occurrences
Repeat Deviation Analysis — Why Did It Come Back?
Dev ID | Dept | Category | Non-Conformance Type | Root Cause Summary | Previous Deviation Ref | Previous CAPA Description | Prior CAPA Effectiveness | Reason Recurrence Occurred
DEV/HR/23/001 | HR | Minor | Documentation / procedural non-compliance | Human Performance › Omission Error › Responsibility boundaries ineffective | Not explicitly cited as a prior deviation reference; recurrence evidenced by review of other PPTs | During investigation, 4 of 5 prepared PPTs (PPT/001-R00, PPT/006-R00, PPT/011-R00, PPT/020-R00) showed the same formatting nonconformance: missing company name and incorrect date format. | ✘ Not Executed | The same formatting error recurred across multiple PPTs, indicating the prior control/training/review process did not detect the omission.
DEV/PK/25/002 | PK | Major | Data logging / temperature monitoring system failure due to HMI time synchronization loss and communication failure | Equipment / Instrument › Electrical Failure | — (no explicit ref) | Previous CAPA identified HMI time synchronization loss due to battery discharge or malfunction; actions included battery replacement for the respective cold room and incorporation of HMI battery verification instructions into the AMC/PM checklist for Blue Star make walk-in cold rooms. | ✘ Not Executed | The CAPA recommendation for battery replacement was not executed because the batteries were considered to be in good condition, allowing recurrence of the synchronization issue.
DEV/PR/24/016 | PR | Minor | Document control / uncontrolled missing controlled copy of approved vendor list (AVL) and related SOP in production department | Human Error › Omission Error › Responsibility boundaries ineffective | — (no explicit ref) | A similar document-control issue was identified for SOP/DP/QA/045-R00 (cleaning validation SOP) being missing from production file handling. | ⚠ Partial | Controlled copies were kept in a common file/table in the production office with no lock-and-key or designated storage, allowing multiple-person use and misplacement.
DEV/PR/25/003 | PR | Minor | Audit trail data not visible / audit trail generation failure in SCADA system | Equipment › Computer system › Scheduled cleanup task failure / database table full (IFIXALARMS) | DEV/QC/25/001 | Mentioned as an incidence at M.J. Biopharm Pune; activities were halted due to that incident. | ✘ Not Executed | The issue recurred because the cleanup task had not been functioning correctly, allowing the IFIXALARMS table to become full again.
DEV/PR/25/004 | PR | Minor | Equipment/Instrument malfunction leading to sterilizer cycle temperature overshoot and cycle abort | Equipment/Instrument › Mechanical failure | DEV/PR/25/002 | Previous deviation reference noted in the report; current report states the issue occurred for the first time in this context, but also flags recurrence by deviation reference. | ✘ Not Executed | No defined regular frequency existed for verification of jacket steam inlet valve (101), allowing the mechanical issue to go undetected.
⚑ Summary: 5 of 12 deviations are repeats (41%). In 4 of 5 cases the prior CAPA was never executed.